## Header
Title | NIRPS transmission budget
Version | 3.2
Date | 2019-02-13
Author | N. Blind | University of Geneva
Notes: | - Seeing values given at 500nm
Version
Changes | 3 | Measured AR coating of FISBA Fiber Link optics
| 3.1 | Updated coupling values with new wf data cubes ad updated RON model in CAOS
Updated with measurements of BBAR ocating for Front-End optics
Removed "IYJHK AR model" based on an old mdoel for the ADC coating (measurments available and added)
| 3.2 | Added dichroic measurements
| 3.3 | BE tranmission updated based on NIRPS-4300-ULA-SP-342-1.0_BE_Throughput_asbuilt (AIT25)
| 3.4 | FL transmission updated based on NIRPS-4200-IAC-TR-160_1.0_fiber_Link_test_report

## Total
Total trhoughput NIRPS-IJH
lbd [nm] | NIRPS + TEL Imag=10 | NIRPS Imag=10 | NIRPS Imag=10+ATM | NIRPS+TEL Imag=11 | NIRPS Imag=11 | NIRPS+TEL Imag=12 | NIRPS Imag=12 | NIRPS + TEL Seeing 0.9" | NIRPS+TEL 100% coupling | WFS channel | YJH channel
700 | [9 empty] | 0.63559 |  | TEL+NIRPS average transmission, seeing=0.9"
710 | [9 empty] | 0.62706 |  | Average | AO I<10 | AO I=12 | AO I=13 | Seeing
720 | [9 empty] | 0.62083 |  | YJH | 0.084955 | 0.067594 | 0.042592 | 0.098132
730 | [9 empty] | 0.61053
740 | [9 empty] | 0.59888
750 | [9 empty] | 0.58741
760 | [9 empty] | 0.57501 |  | WFS average throughput
770 | [9 empty] | 0.56826 |  | Mode | Wavelength | T
780 | [9 empty] | 0.55661 |  | Standard | 700-950 | 0.57761
790 | [9 empty] | 0.55298
800 | [9 empty] | 0.54798
810 | [9 empty] | 0.54616
820 | [9 empty] | 0.54216
830 | [9 empty] | 0.54355
840 | [9 empty] | 0.54381
850 | [9 empty] | 0.54779
860 | [9 empty] | 0.55317
870 | [9 empty] | 0.55908
880 | [9 empty] | 0.56394
890 | [9 empty] | 0.57091
900 | [9 empty] | 0.57643
910 | [9 empty] | 0.58549
920 | [9 empty] | 0.59373
930 | [9 empty] | 0.60167
940 | [9 empty] | 0.60759
950 | [9 empty] | 0.60114
960
970
980 | 0.039728 | 0.051175 | 0.039096 | 0.029448 | 0.037546 | 0.017879 | 0.023072 | 0.05569 | 0.078058 |  | 0.63527
990 | 0.041345 | 0.052879 | 0.040899 | 0.0309 | 0.039064 | 0.018565 | 0.023447 | 0.056973 | 0.079742 |  | 0.64845
1000 | 0.042493 | 0.054248 | 0.042145 | 0.031717 | 0.039788 | 0.019124 | 0.024049 | 0.058778 | 0.082152 |  | 0.65282
1010 | 0.044061 | 0.056061 | 0.043726 | 0.032783 | 0.040826 | 0.019837 | 0.024765 | 0.06086 | 0.084942 |  | 0.6592
1020 | 0.045949 | 0.058272 | 0.045131 | 0.034446 | 0.042603 | 0.020896 | 0.025958 | 0.063339 | 0.088279 |  | 0.6654
1030 | 0.048178 | 0.060936 | 0.047874 | 0.036407 | 0.044722 | 0.02214 | 0.02727 | 0.066078 | 0.091969 |  | 0.67291
1040 | 0.05061 | 0.063712 | 0.050276 | 0.038433 | 0.046914 | 0.023437 | 0.028659 | 0.06894 | 0.095821 |  | 0.67901
1050 | 0.053154 | 0.06654 | 0.052729 | 0.040496 | 0.049156 | 0.024762 | 0.030114 | 0.071879 | 0.099771 |  | 0.684
1060 | 0.055787 | 0.069465 | 0.055017 | 0.042576 | 0.051431 | 0.026129 | 0.031027 | 0.073447 | 0.10181 |  | 0.68862
1070 | 0.057322 | 0.070739 | 0.056714 | 0.043747 | 0.052582 | 0.026839 | 0.032958 | 0.077697 | 0.10756 |  | 0.68031
1080 | 0.060986 | 0.07451 | 0.060388 | 0.046355 | 0.055424 | 0.028521 | 0.034127 | 0.080155 | 0.11082 |  | 0.69873
1090 | 0.063218 | 0.076745 | 0.062826 | 0.048004 | 0.057125 | 0.029619 | 0.035382 | 0.082752 | 0.11426 |  | 0.7026
1100 | 0.065568 | 0.079346 | 0.064669 | 0.049729 | 0.058943 | 0.030851 | 0.036591 | 0.085167 | 0.11744 |  | 0.7081
1110 | 0.067708 | 0.081314 | 0.060937 | 0.051054 | 0.060284 | 0.031844 | 0.03752 | 0.0869 | 0.11967 |  | 0.71124
1120 | 0.069108 | 0.082456 | 0.015653 | 0.051917 | 0.061097 | 0.03227 | 0.036987 | 0.086044 | 0.11841 |  | 0.71211
1130 | 0.068656 | 0.081337 | 0.062663 | 0.051377 | 0.060235 | 0.031823 | 0.038192 | 0.089273 | 0.12276 |  | 0.69618
1140 | 0.071407 | 0.084149 | 0.059939 | 0.05368 | 0.062732 | 0.032897 | 0.03839 | 0.090114 | 0.12383 |  | 0.71496
1150 | 0.072267 | 0.084944 | 1.905e-14 | 0.054697 | 0.06374 | 0.033169 | 0.038584 | 0.090917 | 0.12484 |  | 0.71598
1160 | 0.073122 | 0.085709 | 0.069993 | 0.055694 | 0.064722 | 0.033426 | 0.03877 | 0.091701 | 0.12583 |  | 0.71669
1170 | 0.074159 | 0.086654 | 0.068894 | 0.056666 | 0.065674 | 0.034158 | 0.039294 | 0.09193 | 0.12605 |  | 0.71726
1180 | 0.074749 | 0.086802 | 0.053931 | 0.057114 | 0.066023 | 0.034575 | 0.040045 | 0.092674 | 0.12699 |  | 0.71377
1190 | 0.075724 | 0.087639 | 0.058822 | 0.057984 | 0.066897 | 0.035247 | 0.040736 | 0.093215 | 0.12765 |  | 0.71655
1200 | 0.076614 | 0.088804 | 0.0609 | 0.059102 | 0.068027 | 0.03607 | 0.041169 | 0.093202 | 0.12754 |  | 0.71948
1210 | 0.077065 | 0.089211 | 0.060573 | 0.059754 | 0.068578 | 0.036611 | 0.04216 | 0.094495 | 0.12923 |  | 0.71552
1220 | 0.078459 | 0.090481 | 0.077557 | 0.060933 | 0.069798 | 0.037448 | 0.043073 | 0.095595 | 0.13065 |  | 0.71932
1230 | 0.079796 | 0.091343 | 0.079174 | 0.061857 | 0.070706 | 0.03813 | 0.043661 | 0.095984 | 0.13111 |  | 0.72381
1240 | 0.080534 | 0.092171 | 0.079793 | 0.062717 | 0.071533 | 0.038806 | 0.044318 | 0.096526 | 0.13177 |  | 0.72665
1250 | 0.081316 | 0.093352 | 0.080105 | 0.063756 | 0.072634 | 0.039593 | 0.045197 | 0.097441 | 0.13294 |  | 0.72843
1260 | 0.082536 | 0.094745 | 0.079639 | 0.06427 | 0.073107 | 0.040493 | 0.046153 | 0.098544 | 0.13436 |  | 0.73102
1270 | 0.083701 | 0.095892 | 0.079934 | 0.06486 | 0.073605 | 0.040932 | 0.046658 | 0.10002 | 0.1363 |  | 0.7342
1280 | 0.085063 | 0.097237 | 0.084025 | 0.066049 | 0.074826 | 0.041427 | 0.046931 | 0.10093 | 0.13747 |  | 0.73875
1290 | 0.085979 | 0.098208 | 0.085412 | 0.066965 | 0.075762 | 0.041746 | 0.047195 | 0.1018 | 0.13857 |  | 0.73973
1300 | 0.086876 | 0.098819 | 0.084573 | 0.067286 | 0.076032 | 0.041907 | 0.047604 | 0.10297 | 0.14009 |  | 0.74008
1310 | 0.088061 | 0.099841 | 0.087268 | 0.067896 | 0.076616 | 0.042246 | 0.047719 | 0.10352 | 0.14077 |  | 0.74502
1320 | 0.08879 | 0.10076 | 0.06946 | 0.068941 | 0.077674 | 0.042601 | 0.047889 | 0.104 | 0.14135 |  | 0.74692
1330 | 0.089416 | 0.10166 | 0.015379 | 0.069955 | 0.078723 | 0.042933 | 0.04796 | 0.10423 | 0.14159 |  | 0.74612
1340 | 0.089814 | 0.10222 | 0.013032 | 0.070689 | 0.079493 | 0.043098 | 0.048516 | 0.10546 | 0.1432 |  | 0.74118
1350 | 0.091134 | 0.10285 | 0.014135 | 0.071484 | 0.080321 | 0.043298 | 0.048913 | 0.10636 | 0.14435 |  | 0.74256
1360 | 0.092184 | 0.10324 | 4.1059e-05 | 0.072129 | 0.080957 | 0.04341 | 0.048757 | 0.10608 | 0.14391 |  | 0.74737
1370 | 0.092318 | 0.10296 | 4.8079e-06 | 0.072233 | 0.080972 | 0.043669 | 0.049077 | 0.1057 | 0.14333 |  | 0.74882
1380 | 0.092347 | 0.10307 | 0.00080665 | 0.0726 | 0.081282 | 0.044093 | 0.049559 | 0.10567 | 0.14323 |  | 0.75167
1390 | 0.092658 | 0.1036 | 0.023563 | 0.073241 | 0.08192 | 0.044682 | 0.049865 | 0.10525 | 0.1426 |  | 0.75535
1400 | 0.09265 | 0.10367 | 9.4411e-10 | 0.073548 | 0.082178 | 0.045075 | 0.050113 | 0.10473 | 0.14183 |  | 0.75443
1410 | 0.092557 | 0.10365 | 0.0011218 | 0.073798 | 0.082362 | 0.045432 | 0.050169 | 0.10383 | 0.14056 |  | 0.75154
1420 | 0.091902 | 0.10273 | 0.051676 | 0.073525 | 0.081984 | 0.045028 | 0.050745 | 0.10501 | 0.14211 |  | 0.74427
1430 | 0.09308 | 0.10426 | 0.02986 | 0.074976 | 0.08356 | 0.045678 | 0.051016 | 0.10552 | 0.14274 |  | 0.75261
1440 | 0.093713 | 0.10496 | 0.00063978 | 0.075078 | 0.083622 | 0.045973 | 0.051024 | 0.10549 | 0.14266 |  | 0.75469
1450 | 0.0939 | 0.1051 | 0.055636 | 0.074811 | 0.083237 | 0.046042 | 0.050638 | 0.1047 | 0.14154 |  | 0.75281
1460 | 0.093374 | 0.10424 | 0.054614 | 0.074579 | 0.082883 | 0.045675 | 0.050889 | 0.10523 | 0.14221 |  | 0.745
1470 | 0.094127 | 0.10447 | 0.05914 | 0.074996 | 0.083277 | 0.046108 | 0.051612 | 0.10578 | 0.14291 |  | 0.74751
1480 | 0.094913 | 0.10462 | 0.081208 | 0.075317 | 0.083575 | 0.046499 | 0.051628 | 0.10487 | 0.14163 |  | 0.75405
1490 | 0.094406 | 0.10419 | 3.6422e-07 | 0.075213 | 0.0834 | 0.046626 | 0.051738 | 0.10417 | 0.14064 |  | 0.75505
1500 | 0.094035 | 0.10461 | 0.092014 | 0.075678 | 0.083863 | 0.047134 | 0.05225 | 0.10427 | 0.14074 |  | 0.75606
1510 | 0.094393 | 0.10527 | 0.093619 | 0.076304 | 0.084508 | 0.047743 | 0.052862 | 0.10456 | 0.1411 |  | 0.75675
1520 | 0.094823 | 0.10624 | 0.094358 | 0.077207 | 0.085455 | 0.048456 | 0.053734 | 0.10563 | 0.14251 |  | 0.75703
1530 | 0.09594 | 0.10699 | 0.092265 | 0.077931 | 0.086218 | 0.049058 | 0.054304 | 0.1061 | 0.14309 |  | 0.7588
1540 | 0.097287 | 0.10771 | 0.097073 | 0.078529 | 0.086828 | 0.04966 | 0.054215 | 0.10528 | 0.14195 |  | 0.75967
1550 | 0.095973 | 0.10758 | 0.095253 | 0.078534 | 0.086757 | 0.049887 | 0.055598 | 0.10652 | 0.1447 |  | 0.75074
1560 | 0.097251 | 0.10819 | 0.097115 | 0.078474 | 0.086634 | 0.050442 | 0.055969 | 0.10744 | 0.14477 |  | 0.75788
1570 | 0.098094 | 0.10937 | 0.089609 | 0.078942 | 0.087107 | 0.051049 | 0.056384 | 0.10822 | 0.14577 |  | 0.7621
1580 | 0.098809 | 0.11016 | 0.096774 | 0.079943 | 0.088157 | 0.05148 | 0.056655 | 0.10874 | 0.14641 |  | 0.76331
1590 | 0.099274 | 0.11039 | 0.099145 | 0.080542 | 0.088762 | 0.051651 | 0.056649 | 0.10872 | 0.14634 |  | 0.76224
1600 | 0.099222 | 0.1102 | 0.098428 | 0.080795 | 0.089013 | 0.051609 | 0.057437 | 0.11019 | 0.14827 |  | 0.75881
1610 | 0.10059 | 0.11151 | 0.09965 | 0.082159 | 0.090482 | 0.052274 | 0.057795 | 0.11086 | 0.14911 |  | 0.76682
1620 | 0.1014 | 0.11228 | 0.10121 | 0.082977 | 0.091333 | 0.053049 | 0.058527 | 0.11129 | 0.14964 |  | 0.77024
1630 | 0.10195 | 0.11313 | 0.10179 | 0.083831 | 0.09225 | 0.053849 | 0.059283 | 0.11175 | 0.15018 |  | 0.77223
1640 | 0.10256 | 0.1137 | 0.10233 | 0.084459 | 0.092918 | 0.054516 | 0.059909 | 0.11195 | 0.15039 |  | 0.77276
1650 | 0.10296 | 0.11434 | 0.10268 | 0.085167 | 0.093653 | 0.055238 | 0.060615 | 0.11233 | 0.15082 |  | 0.77224
1660 | 0.10351 | 0.1146 | 0.10264 | 0.085602 | 0.094056 | 0.055791 | 0.061003 | 0.11216 | 0.15052 |  | 0.77125
1670 | 0.10349 | 0.11434 | 0.10321 | 0.08566 | 0.094065 | 0.056073 | 0.062138 | 0.1134 | 0.15209 |  | 0.76795
1680 | 0.10481 | 0.11551 | 0.10406 | 0.086808 | 0.095237 | 0.057074 | 0.062898 | 0.11396 | 0.15277 |  | 0.77568
1690 | 0.10544 | 0.11586 | 0.10401 | 0.087303 | 0.09573 | 0.05765 | 0.063231 | 0.1137 | 0.15236 |  | 0.77957
1700 | 0.10531 | 0.11547 | 0.10496 | 0.087194 | 0.095607 | 0.057832 | 0.063206 | 0.1128 | 0.15102 |  | 0.7798
1710 | 0.10468 | 0.11456 | 0.1041 | 0.086748 | 0.095053 | 0.057787 | 0.062764 | 0.11122 | 0.14882 |  | 0.77778
1720 | 0.10315 | 0.11276 | 0.029727 | 0.084795 | 0.092851 | 0.05673 | 0.062328 | 0.11092 | 0.14835 |  | 0.77148
1730 | 0.10276 | 0.11217 | 0.10143 | 0.083736 | 0.091674 | 0.056265 | 0.061882 | 0.11058 | 0.14778 |  | 0.77423
1740 | 0.10239 | 0.11162 | 0.096053 | 0.083645 | 0.091564 | 0.055812 | 0.061155 | 0.10973 | 0.14653 |  | 0.77788
1750 | 0.10154 | 0.11059 | 0.056577 | 0.083193 | 0.091053 | 0.055125 | 0.060187 | 0.10842 | 0.14469 |  | 0.77878
1760 | 0.10025 | 0.10881 | 0.098564 | 0.082123 | 0.08989 | 0.054052 | 0.059047 | 0.1068 | 0.14239 |  | 0.77682
1770 | 0.098765 | 0.10699 | 0.097155 | 0.0809 | 0.088578 | 0.053438 | 0.05827 | 0.1047 | 0.13948 |  | 0.77505
1780 | 0.096887 | 0.10471 | 0.067675 | 0.079319 | 0.086864 | 0.052591 | 0.056963 | 0.10169 | 0.13537 |  | 0.77173
1790 | 0.094154 | 0.10148 | 0.056746 | 0.076999 | 0.084343 | 0.051244 | 0.05451 | 0.096703 | 0.12861 |  | 0.76311
1800 | 0.094154 | 0.10148 | 0.056746 | 0.076999 | 0.084343 | 0.051244 | 0.05451 | 0.096703 | 0.12861 |  | 0.74107
1810 | [10 empty] | 0.74129
1820 | [10 empty] | 0
1830 | [10 empty] | 0
1840 | [10 empty] | 0
1850 | [10 empty] | 0
1860 | [10 empty] | 0
1870 | [10 empty] | 0
1880 | [10 empty] | 0
1890 | [10 empty] | 0
1900 | [10 empty] | 0
1910 | [10 empty] | 0
1920 | [10 empty] | 0
1930 | [10 empty] | 0
1940 | [10 empty] | 0
1950 | [10 empty] | 0
1960 | [10 empty] | 0
1970 | [10 empty] | 0
1980 | [10 empty] | 0
1990 | [10 empty] | 0
2000 | [10 empty] | 0
2010 | [10 empty] | 0
2020 | [10 empty] | 0
2030 | [10 empty] | 0
2040 | [10 empty] | 0
2050 | [10 empty] | 0
2060 | [10 empty] | 0
2070 | [10 empty] | 0
2080 | [10 empty] | 0
2090 | [10 empty] | 0
2100 | [10 empty] | 0
2110 | [10 empty] | 0
2120 | [10 empty] | 0
2130 | [10 empty] | 0
2140 | [10 empty] | 0
2150 | [10 empty] | 0
2160 | [10 empty] | 0
2170 | [10 empty] | 0
2180 | [10 empty] | 0
2190 | [10 empty] | 0
2200 | [10 empty] | 0
2210 | [10 empty] | 0
2220 | [10 empty] | 0
2230 | [10 empty] | 0
2240 | [10 empty] | 0
2250 | [10 empty] | 0
2260 | [10 empty] | 0
2270 | [10 empty] | 0
2280 | [10 empty] | 0
2290 | [10 empty] | 0
2300 | [10 empty] | 0
2310 | [10 empty] | 0
2320 | [10 empty] | 0
2330 | [10 empty] | 0
2340 | [10 empty] | 0
2350 | [10 empty] | 0
2360 | [10 empty] | 0
2370 | [10 empty] | 0
2380 | [10 empty] | 0
2390 | [10 empty] | 0
2400 | [10 empty] | 0

## Atm.
ATMOSPHERIC TRANSMITTANCE FROM TAPAS MODEL
lbd [nm] | Thoughput
381 | 0.7148
391 | 0.7407
401 | 0.7632
411 | 0.7835
421 | 0.8017
431 | 0.8181
441 | 0.8323
451 | 0.8457
461 | 0.8566
471 | 0.8664
481 | 0.8731
491 | 0.8832
501 | 0.8876
511 | 0.8933
521 | 0.8974
531 | 0.8961
541 | 0.8994
551 | 0.9016
561 | 0.9009
571 | 0.8959
581 | 0.9017
591 | 0.8938
601 | 0.9058
611 | 0.9146
621 | 0.9212
631 | 0.9223
641 | 0.9362
651 | 0.943
661 | 0.9484
671 | 0.9541
681 | 0.9585
691 | 0.9554
701 | 0.954
711 | 0.9653
721 | 0.9496
731 | 0.9539
741 | 0.9741
751 | 0.9762
761 | 0.6165
771 | 0.9777
781 | 0.9797
791 | 0.9798
801 | 0.9821
811 | 0.9765
821 | 0.9682
831 | 0.9758
841 | 0.9834
851 | 0.9855
861 | 0.9864
871 | 0.9878
881 | 0.9886
891 | 0.9888
901 | 0.9756
911 | 0.9811
921 | 0.9803
931 | 0.6565
941 | 0.42
951 | 0.9423
961 | 0.05383
971 | 0.983
981 | 0.9841
991 | 0.9892
1001 | 0.9918
1011 | 0.9924
1021 | 0.9822
1031 | 0.9937
1041 | 0.9934
1051 | 0.992
1061 | 0.9862
1071 | 0.9894
1081 | 0.9902
1091 | 0.9938
1101 | 0.9863
1111 | 0.9
1121 | 0.2265
1131 | 0.9127
1141 | 0.8394
1151 | 2.636e-13
1161 | 0.9572
1171 | 0.929
1181 | 0.7215
1191 | 0.7768
1201 | 0.7949
1211 | 0.786
1221 | 0.9885
1231 | 0.9922
1241 | 0.9908
1251 | 0.9851
1261 | 0.9649
1271 | 0.955
1281 | 0.9878
1291 | 0.9934
1301 | 0.9735
1311 | 0.991
1321 | 0.7823
1331 | 0.172
1341 | 0.1451
1351 | 0.1551
1361 | 0.0004454
1371 | 5.208e-05
1381 | 0.008735
1391 | 0.2543
1401 | 1.019e-08
1411 | 0.01212
1421 | 0.5623
1431 | 0.3208
1441 | 0.006827
1451 | 0.5925
1461 | 0.5849
1471 | 0.6283
1481 | 0.8556
1491 | 3.858e-06
1501 | 0.9785
1511 | 0.9918
1521 | 0.9951
1531 | 0.9617
1541 | 0.9978
1551 | 0.9925
1561 | 0.9986
1571 | 0.9135
1581 | 0.9794
1591 | 0.9987
1601 | 0.992
1611 | 0.9907
1621 | 0.9981
1631 | 0.9984
1641 | 0.9978
1651 | 0.9973
1661 | 0.9916
1671 | 0.9973
1681 | 0.9928
1691 | 0.9865
1701 | 0.9967
1711 | 0.9945
1721 | 0.2882
1731 | 0.987
1741 | 0.9381
1751 | 0.5572
1761 | 0.9832
1771 | 0.9837
1781 | 0.6985
1791 | 0.6027
1801 | 0.3364
1811 | 0.3117
1821 | -4.248e-18
1831 | 3.227e-06
1841 | -3.976e-28
1851 | 0.07733
1861 | 0.02697
1871 | -1.484e-47
1881 | 0.00843
1891 | 1.698e-09
1901 | 0.01302
1911 | 2.717e-13
1921 | 0.008408
1931 | 0.01636
1941 | 0.0004959
1951 | 0.7635
1961 | 0.173
1971 | 0.003476
1981 | 0.7625
1991 | 0.9773
2001 | 1.429e-09
2011 | 0.001756
2021 | 0.5175
2031 | 0.9849
2041 | 0.8612
2051 | 0.4863
2061 | 0.8414
2071 | 0.9396
2081 | 0.9823
2091 | 0.9755
2101 | 0.9278
2111 | 0.9901
2121 | 0.996
2131 | 0.9988
2141 | 0.9988
2151 | 0.977
2161 | 0.9956
2171 | 0.998
2181 | 0.9982
2191 | 0.9984
2201 | 0.972
2211 | 0.9959
2221 | 0.852
2231 | 0.6461
2241 | 0.9858
2251 | 0.9603
2261 | 0.9685
2271 | 0.9706
2281 | 0.9685
2291 | 0.9315
2301 | 0.9819
2311 | 0.9477
2321 | 0.9679
2331 | 0.757
2341 | 0.9529
2351 | 0.0337
2361 | 0.2929
2371 | 0.1149
2381 | 0.9722
2391 | 0.4594
2401 | 0.9539

## Telescope
Telescope made of 2 Bare Al coated mirrors
| Clean telescope | From Thorlabs proteccted Al coating. At 670nm, R=0.915, against R=0.88 measured at NTT. Values are scaled in consequence.
| Dirty telescope | From VLT-TRE-ESO-11492-1043 "Dust Survey", NTT M1 coating reflectance degraded from 88% (fresh coating) to 72% after 4 years (1992-1996). Measurement at 670nm. Only M1 affected. Dirty telescope is also scaled to match R=88% at 670 when clean (not the case of Thorlabs data).
LBD [mum] | Telescope clean | Telescope dirty
0.38 | 0.80649 | 0.72307
0.39 | 0.83414 | 0.74786
0.4 | 0.85168 | 0.76359
0.41 | 0.86325 | 0.77396
0.42 | 0.87111 | 0.781
0.43 | 0.8761 | 0.78548
0.44 | 0.87956 | 0.78858
0.45 | 0.88154 | 0.79036
0.46 | 0.88322 | 0.79186
0.47 | 0.88393 | 0.7925
0.48 | 0.88404 | 0.7926
0.49 | 0.88316 | 0.79181
0.5 | 0.88214 | 0.79089
0.51 | 0.88056 | 0.78948
0.52 | 0.87862 | 0.78774
0.53 | 0.87614 | 0.78552
0.54 | 0.87341 | 0.78307
0.55 | 0.87025 | 0.78023
0.56 | 0.86438 | 0.77497
0.57 | 0.86068 | 0.77166
0.58 | 0.85652 | 0.76793
0.59 | 0.85206 | 0.76393
0.6 | 0.84715 | 0.75953
0.61 | 0.84205 | 0.75495
0.62 | 0.8366 | 0.75006
0.63 | 0.83056 | 0.74465
0.64 | 0.82417 | 0.73892
0.65 | 0.81745 | 0.7329
0.66 | 0.81053 | 0.72669
0.67 | 0.80307 | 0.72
0.68 | 0.79521 | 0.71295
0.69 | 0.78704 | 0.70563
0.7 | 0.7784 | 0.69788
0.71 | 0.76931 | 0.68974
0.72 | 0.76005 | 0.68143
0.73 | 0.75016 | 0.67256
0.74 | 0.73988 | 0.66335
0.75 | 0.72961 | 0.65414
0.76 | 0.71928 | 0.64488
0.77 | 0.70869 | 0.63539
0.78 | 0.69891 | 0.62662
0.79 | 0.68956 | 0.61823
0.8 | 0.682 | 0.61146
0.81 | 0.67586 | 0.60595
0.82 | 0.67219 | 0.60267
0.83 | 0.67114 | 0.60172
0.84 | 0.67292 | 0.60332
0.85 | 0.67717 | 0.60713
0.86 | 0.68361 | 0.6129
0.87 | 0.6918 | 0.62024
0.88 | 0.70116 | 0.62864
0.89 | 0.70314 | 0.63041
0.9 | 0.71352 | 0.63972
0.91 | 0.72343 | 0.6486
0.92 | 0.7334 | 0.65754
0.93 | 0.7431 | 0.66624
0.94 | 0.75212 | 0.67432
0.95 | 0.7607 | 0.68202
0.96 | 0.76878 | 0.68926
0.97 | 0.77679 | 0.69644
0.98 | 0.7843 | 0.70318
0.99 | 0.79102 | 0.7092
1 | 0.79713 | 0.71468
1.01 | 0.80298 | 0.71992
1.02 | 0.80854 | 0.72491
1.03 | 0.81408 | 0.72987
1.04 | 0.81922 | 0.73449
1.05 | 0.82381 | 0.7386
1.06 | 0.82784 | 0.74221
1.07 | 0.83198 | 0.74592
1.08 | 0.83639 | 0.74987
1.09 | 0.84033 | 0.75341
1.1 | 0.84369 | 0.75642
1.11 | 0.84689 | 0.75929
1.12 | 0.84976 | 0.76186
1.13 | 0.85295 | 0.76473
1.14 | 0.8557 | 0.76719
1.15 | 0.85813 | 0.76937
1.16 | 0.86051 | 0.7715
1.17 | 0.86283 | 0.77359
1.18 | 0.86506 | 0.77558
1.19 | 0.86676 | 0.77711
1.2 | 0.8688 | 0.77893
1.21 | 0.87133 | 0.7812
1.22 | 0.87299 | 0.7827
1.23 | 0.87485 | 0.78436
1.24 | 0.87676 | 0.78607
1.25 | 0.87776 | 0.78697
1.26 | 0.87912 | 0.78819
1.27 | 0.88119 | 0.79005
1.28 | 0.8827 | 0.7914
1.29 | 0.88389 | 0.79246
1.3 | 0.88497 | 0.79344
1.31 | 0.88619 | 0.79453
1.32 | 0.88757 | 0.79576
1.33 | 0.88863 | 0.79671
1.34 | 0.88925 | 0.79727
1.35 | 0.88998 | 0.79792
1.36 | 0.89095 | 0.79879
1.37 | 0.89207 | 0.7998
1.38 | 0.89319 | 0.8008
1.39 | 0.89406 | 0.80158
1.4 | 0.89498 | 0.80241
1.41 | 0.89602 | 0.80334
1.42 | 0.89682 | 0.80406
1.43 | 0.89727 | 0.80446
1.44 | 0.89783 | 0.80496
1.45 | 0.89877 | 0.8058
1.46 | 0.89981 | 0.80674
1.47 | 0.90056 | 0.80741
1.48 | 0.90118 | 0.80797
1.49 | 0.90184 | 0.80856
1.5 | 0.9024 | 0.80906
1.51 | 0.90292 | 0.80952
1.52 | 0.90349 | 0.81003
1.53 | 0.90388 | 0.81039
1.54 | 0.90443 | 0.81088
1.55 | 0.90522 | 0.81158
1.56 | 0.90581 | 0.81212
1.57 | 0.90627 | 0.81253
1.58 | 0.90682 | 0.81302
1.59 | 0.90739 | 0.81353
1.6 | 0.90768 | 0.81379
1.61 | 0.90802 | 0.8141
1.62 | 0.90852 | 0.81454
1.63 | 0.90874 | 0.81474
1.64 | 0.90895 | 0.81494
1.65 | 0.90938 | 0.81532
1.66 | 0.91012 | 0.81598
1.67 | 0.91065 | 0.81645
1.68 | 0.91148 | 0.81721
1.69 | 0.91197 | 0.81764
1.7 | 0.912 | 0.81767
1.71 | 0.91262 | 0.81822
1.72 | 0.91324 | 0.81878
1.73 | 0.91342 | 0.81894
1.74 | 0.91351 | 0.81902
1.75 | 0.91368 | 0.81917
1.76 | 0.91359 | 0.8191
1.77 | 0.91331 | 0.81884
1.78 | 0.91313 | 0.81868
1.79 | 0.91293 | 0.8185
1.8 | 0.91294 | 0.81851
1.81 | 0.9132 | 0.81874
1.82 | 0.91387 | 0.81934
1.83 | 0.91451 | 0.81992
1.84 | 0.91489 | 0.82026
1.85 | 0.91515 | 0.82049
1.86 | 0.91541 | 0.82073
1.87 | 0.9154 | 0.82071
1.88 | 0.91521 | 0.82055
1.89 | 0.91489 | 0.82026
1.9 | 0.91483 | 0.8202
1.91 | 0.91483 | 0.8202
1.92 | 0.91478 | 0.82016
1.93 | 0.91482 | 0.82019
1.94 | 0.915 | 0.82035
1.95 | 0.91534 | 0.82066
1.96 | 0.91583 | 0.8211
1.97 | 0.91635 | 0.82157
1.98 | 0.91688 | 0.82204
1.99 | 0.9172 | 0.82233
2 | 0.91725 | 0.82238
2.01 | 0.91716 | 0.82229
2.02 | 0.91683 | 0.822
2.03 | 0.9164 | 0.82161
2.04 | 0.91597 | 0.82123
2.05 | 0.9157 | 0.82099
2.06 | 0.91556 | 0.82086
2.07 | 0.91521 | 0.82054
2.08 | 0.91515 | 0.82049
2.09 | 0.91561 | 0.8209
2.1 | 0.91605 | 0.8213
2.11 | 0.91679 | 0.82196
2.12 | 0.91729 | 0.82241
2.13 | 0.91789 | 0.82295
2.14 | 0.91845 | 0.82345
2.15 | 0.91861 | 0.8236
2.16 | 0.91878 | 0.82374
2.17 | 0.91843 | 0.82343
2.18 | 0.91824 | 0.82326
2.19 | 0.91816 | 0.82319
2.2 | 0.91763 | 0.82271
2.21 | 0.91723 | 0.82235
2.22 | 0.91692 | 0.82208
2.23 | 0.91666 | 0.82185
2.24 | 0.91642 | 0.82163
2.25 | 0.91631 | 0.82153
2.26 | 0.9165 | 0.8217
2.27 | 0.91684 | 0.82201
2.28 | 0.91697 | 0.82212
2.29 | 0.91761 | 0.8227
2.3 | 0.9182 | 0.82323
2.31 | 0.91936 | 0.82427
2.32 | 0.91985 | 0.8247
2.33 | 0.91964 | 0.82451
2.34 | 0.91985 | 0.8247
2.35 | 0.91986 | 0.82471
2.36 | 0.92006 | 0.82489
2.37 | 0.91923 | 0.82415
2.38 | 0.91908 | 0.82402
2.39 | 0.91885 | 0.82381
2.4 | 0.91839 | 0.82339

## FE YJH
Front end - IJHK channel, until fiber
Optical surfaces | [6 empty] | Type
4 refractive surfaces VIJHK |  |  |  | ADC surfaces |  |  | MEASURED
5 refractive surfaces I |  |  |  | Optimax BBAR YJH (L1+L2 doublets + WFS dichro backside) |  |  | MEASURED
2 Ag mirrors |  |  |  | M3 fold (thorlabs data) + DM241 (ALPAO measurement) |  |  | MEASURED
2 broadband dichroic |  |  |  | R or T > 97% |  |  | SPEC
ADC transmission (from internal refleections) |  |  |  | 95.5% |  |  | SPEC
lbd [nm] | NIR/VIS dichro | YJH/WFS dichro | Front-End /w mirror | Front-End /from FP1
980 | 0.99439 | 0.97989 | 0.80997 | 0.81455
990 | 0.99832 | 0.98921 | 0.81977 | 0.82115
1000 | 0.99829 | 0.98911 | 0.81896 | 0.82036
1010 | 0.99936 | 0.98921 | 0.82094 | 0.82147
1020 | 0.99943 | 0.98789 | 0.82297 | 0.82344
1030 | 0.99976 | 0.98845 | 0.82659 | 0.8268
1040 | 0.99952 | 0.98641 | 0.82885 | 0.82925
1050 | 0.99919 | 0.98561 | 0.83029 | 0.83096
1060 | 0.99876 | 0.98381 | 0.83183 | 0.83286
1070 | 0.97516 | 0.98653 | 0.8177 | 0.83852
1080 | 0.9946 | 0.98629 | 0.83542 | 0.83995
1090 | 0.99367 | 0.9869 | 0.83609 | 0.84142
1100 | 0.99706 | 0.98786 | 0.8393 | 0.84177
1110 | 0.99785 | 0.9867 | 0.83983 | 0.84163
1120 | 0.99719 | 0.98736 | 0.83802 | 0.84038
1130 | 0.97426 | 0.98563 | 0.81621 | 0.83777
1140 | 0.9969 | 0.98744 | 0.83552 | 0.83812
1150 | 0.99801 | 0.98688 | 0.83435 | 0.83601
1160 | 0.99727 | 0.98773 | 0.83287 | 0.83515
1170 | 0.99631 | 0.98837 | 0.83129 | 0.83437
1180 | 0.99133 | 0.98805 | 0.82512 | 0.83233 | FE YJH average | 0.83637
1190 | 0.9925 | 0.98929 | 0.8267 | 0.83295
1200 | 0.99317 | 0.9896 | 0.82814 | 0.83383
1210 | 0.98541 | 0.98857 | 0.82119 | 0.83334
1220 | 0.9885 | 0.98918 | 0.82397 | 0.83356
1230 | 0.9908 | 0.98889 | 0.82736 | 0.83504
1240 | 0.9906 | 0.98944 | 0.82879 | 0.83666
1250 | 0.99099 | 0.98891 | 0.82987 | 0.83742
1260 | 0.99065 | 0.98968 | 0.83153 | 0.83938
1270 | 0.99286 | 0.98706 | 0.83319 | 0.83919
1280 | 0.99383 | 0.98805 | 0.83693 | 0.84212
1290 | 0.99212 | 0.98775 | 0.8369 | 0.84355
1300 | 0.99207 | 0.98683 | 0.83627 | 0.84296
1310 | 0.99426 | 0.98808 | 0.8407 | 0.84555
1320 | 0.99424 | 0.98758 | 0.84154 | 0.84641
1330 | 0.99092 | 0.98829 | 0.83963 | 0.84733
1340 | 0.98379 | 0.98709 | 0.83348 | 0.84721
1350 | 0.98476 | 0.98711 | 0.83435 | 0.84726
1360 | 0.98797 | 0.98893 | 0.83884 | 0.84906
1370 | 0.98901 | 0.98767 | 0.83942 | 0.84875
1380 | 0.99251 | 0.98799 | 0.84156 | 0.84791
1390 | 0.9951 | 0.989 | 0.84485 | 0.84902
1400 | 0.99491 | 0.988 | 0.84296 | 0.84727
1410 | 0.98975 | 0.98798 | 0.83876 | 0.84745
1420 | 0.98122 | 0.98773 | 0.8299 | 0.84578
1430 | 0.9921 | 0.98761 | 0.83877 | 0.84545
1440 | 0.99461 | 0.98899 | 0.84056 | 0.84512
1450 | 0.99234 | 0.98864 | 0.8376 | 0.84407
1460 | 0.9829 | 0.98744 | 0.82796 | 0.84236
1470 | 0.98616 | 0.98805 | 0.83004 | 0.84169
1480 | 0.99354 | 0.98815 | 0.83674 | 0.84218
1490 | 0.99684 | 0.98739 | 0.83723 | 0.83989
1500 | 0.99846 | 0.98759 | 0.83783 | 0.83912
1510 | 0.99908 | 0.98761 | 0.83812 | 0.83889
1520 | 0.99922 | 0.98752 | 0.8379 | 0.83855
1530 | 0.99899 | 0.98885 | 0.83949 | 0.84034
1540 | 0.99756 | 0.98984 | 0.83995 | 0.842
1550 | 0.98619 | 0.98915 | 0.82935 | 0.84097
1560 | 0.99449 | 0.98904 | 0.83668 | 0.84132
1570 | 0.99791 | 0.98974 | 0.84093 | 0.84269
1580 | 0.99813 | 0.98887 | 0.84175 | 0.84333
1590 | 0.99658 | 0.98726 | 0.84004 | 0.84292
1600 | 0.99018 | 0.9876 | 0.83599 | 0.84428
1610 | 0.99597 | 0.98935 | 0.8445 | 0.84791
1620 | 0.99848 | 0.99023 | 0.8478 | 0.8491
1630 | 0.99895 | 0.98996 | 0.84978 | 0.85068
1640 | 0.99887 | 0.98914 | 0.85017 | 0.85113
1650 | 0.99812 | 0.98785 | 0.8492 | 0.85079
1660 | 0.99494 | 0.98697 | 0.84742 | 0.85173
1670 | 0.98806 | 0.98779 | 0.8433 | 0.85349
1680 | 0.99359 | 0.98972 | 0.85101 | 0.8565
1690 | 0.99595 | 0.99078 | 0.85482 | 0.8583
1700 | 0.9957 | 0.9904 | 0.85504 | 0.85874
1710 | 0.99264 | 0.98953 | 0.85225 | 0.85858
1720 | 0.98535 | 0.98886 | 0.84477 | 0.85733
1730 | 0.99034 | 0.9885 | 0.84762 | 0.85589
1740 | 0.9955 | 0.98874 | 0.85153 | 0.85538
1750 | 0.99703 | 0.98957 | 0.85235 | 0.8549
1760 | 0.9973 | 0.99024 | 0.85029 | 0.8526
1770 | 0.99679 | 0.99019 | 0.84861 | 0.85135
1780 | 0.99504 | 0.98949 | 0.84514 | 0.84936
1790 | 0.98975 | 0.98773 | 0.83589 | 0.84455
1800 | 0.96954 | 0.98426 | 0.81174 | 0.83725
1810 |  |  | 0.81174 | 0.83725
1820 |  |  | 0 | 0
1830 |  |  | 0 | 0
1840 |  |  | 0 | 0
1850 |  |  | 0 | 0
1860 |  |  | 0 | 0
1870 |  |  | 0 | 0
1880 |  |  | 0 | 0
1890 |  |  | 0 | 0
1900 |  |  | 0 | 0
1910 |  |  | 0 | 0
1920 |  |  | 0 | 0
1930 |  |  | 0 | 0
1940 |  |  | 0 | 0
1950 |  |  | 0 | 0
1960 |  |  | 0 | 0
1970 |  |  | 0 | 0
1980 |  |  | 0 | 0
1990 |  |  | 0 | 0
2000 |  |  | 0 | 0
2010 |  |  | 0 | 0
2020 |  |  | 0 | 0
2030 |  |  | 0 | 0
2040 |  |  | 0 | 0
2050 |  |  | 0 | 0
2060 |  |  | 0 | 0
2070 |  |  | 0 | 0
2080 |  |  | 0 | 0
2090 |  |  | 0 | 0
2100 |  |  | 0 | 0
2110 |  |  | 0 | 0
2120 |  |  | 0 | 0
2130 |  |  | 0 | 0
2140 |  |  | 0 | 0
2150 |  |  | 0 | 0
2160 |  |  | 0 | 0
2170 |  |  | 0 | 0
2180 |  |  | 0 | 0
2190 |  |  | 0 | 0
2200 |  |  | 0 | 0
2210 |  |  | 0 | 0
2220 |  |  | 0 | 0
2230 |  |  | 0 | 0
2240 |  |  | 0 | 0
2250 |  |  | 0 | 0
2260 |  |  | 0 | 0
2270 |  |  | 0 | 0
2280 |  |  | 0 | 0
2290 |  |  | 0 | 0
2300 |  |  | 0 | 0
2310 |  |  | 0 | 0
2320 |  |  | 0 | 0
2330 |  |  | 0 | 0
2340 |  |  | 0 | 0
2350 |  |  | 0 | 0
2360 |  |  | 0 | 0
2370 |  |  | 0 | 0
2380 |  |  | 0 | 0
2390 |  |  | 0 | 0
2400 |  |  | 0 | 0

## FE WFS
Front end - IJHK channel, until fiber
Optical surfaces
4 refractive surfaces VIJHK |  |  | Optimax BBAR ADC
8 refractive surfaces |  |  | Optimax BBAR WFS (L1, L2, L3, lenslet array)
2 Ag mirrors |  |  | M3 fold + M4 fold (Thorlabs data) + DM241 (ALPAO data)
2 broadband dichroic |  |  | R or T > 97%
lensleL fill factor |  |  | 0.98
ADC transmission (from internal refleections) |  |  | 95.5%
lbd [nm] | NIR/VIS dichro | WFS/YJH dichro | Front-End /w mirror | Front-End /wo dichro /wo M3 mirror2 |  |  | 1: lbd | 2: t | 3: r
700 | 0.99646 | 0.99989 | 0.81654 | 0.85901 | 0.8465 | 0 | 690 | 0.43758 | 0.56242
710 | 0.9893 | 0.99903 | 0.81509 | 0.8654 |  | 1 | 700 | 0.0035437 | 0.99646
720 | 0.99488 | 0.99988 | 0.81684 | 0.86343 |  | 2 | 710 | 0.010696 | 0.9893
730 | 0.99818 | 0.99991 | 0.81387 | 0.85699 |  | 3 | 720 | 0.005118 | 0.99488
740 | 0.99857 | 0.99994 | 0.80942 | 0.85171 |  | 4 | 730 | 0.0018192 | 0.99818
750 | 0.99634 | 0.99998 | 0.8051 | 0.84826 |  | 5 | 740 | 0.0014309 | 0.99857
760 | 0.99213 | 0.99989 | 0.79942 | 0.8455 |  | 6 | 750 | 0.003663 | 0.99634
770 | 0.99522 | 0.99998 | 0.80185 | 0.84503 |  | 7 | 760 | 0.0078652 | 0.99213
780 | 0.98813 | 1 | 0.79639 | 0.84486 |  | 8 | 770 | 0.0047757 | 0.99522
790 | 0.99404 | 0.99999 | 0.80194 | 0.84498 |  | 9 | 780 | 0.011874 | 0.98813
800 | 0.99406 | 1 | 0.80348 | 0.84609 |  | 10 | 790 | 0.0059571 | 0.99404
810 | 0.99055 | 0.99998 | 0.80811 | 0.85175 |  | 11 | 800 | 0.0059393 | 0.99406
820 | 0.98975 | 0.99977 | 0.80656 | 0.84934 |  | 12 | 810 | 0.0094531 | 0.99055
830 | 0.99361 | 0.99999 | 0.80989 | 0.84937 |  | 13 | 820 | 0.01025 | 0.98975
840 | 0.99174 | 0.99983 | 0.80814 | 0.84908 |  | 14 | 830 | 0.0063941 | 0.99361
850 | 0.99192 | 0.99991 | 0.80893 | 0.84845 |  | 15 | 840 | 0.0082645 | 0.99174
860 | 0.99105 | 0.99986 | 0.80919 | 0.84777 |  | 16 | 850 | 0.008076 | 0.99192
870 | 0.99237 | 0.99994 | 0.80816 | 0.84529 |  | 17 | 860 | 0.0089497 | 0.99105
880 | 0.99231 | 0.99618 | 0.80429 | 0.84022 |  | 18 | 870 | 0.0076288 | 0.99237
890 | 0.99679 | 0.99999 | 0.81195 | 0.84401 |  | 19 | 880 | 0.0076936 | 0.99231
900 | 0.99338 | 0.99997 | 0.80786 | 0.84182 |  | 20 | 890 | 0.0032055 | 0.99679
910 | 0.99477 | 0.99999 | 0.80933 | 0.8419 |  | 21 | 900 | 0.0066155 | 0.99338
920 | 0.99774 | 0.99996 | 0.80956 | 0.83948 |  | 22 | 910 | 0.0052271 | 0.99477
930 | 0.999 | 0.99998 | 0.80967 | 0.83802 |  | 23 | 920 | 0.0022648 | 0.99774
940 | 0.99952 | 0.99978 | 0.80784 | 0.83505 |  | 24 | 930 | 0.00099616 | 0.999
950 | 0.99957 | 0.9804 | 0.79025 | 0.81632 |  | 25 | 940 | 0.00047565 | 0.99952
960 | 0.99861 | 0.93818 | 0.75298 | 0.7783 |  | 26 | 950 | 0.00042566 | 0.99957
970 | 0.99189 | 0.52662 | 0.41814 | 0.43504 |  | 27 | 960 | 0.0013919 | 0.99861
980 | 0.99439 | 0.020114 | 0.01594 | 0.016537 |  | 28 | 970 | 0.0081146 | 0.99189
[6 empty] | 29 | 980 | 0.005613 | 0.99439
[6 empty] | 30 | 990 | 0.0016776 | 0.99832
[6 empty] | 31 | 1000 | 0.001706 | 0.99829
[6 empty] | 32 | 1010 | 0.00064257 | 0.99936
[6 empty] | 33 | 1020 | 0.00057272 | 0.99943
[6 empty] | 34 | 1030 | 0.00024244 | 0.99976
[6 empty] | 35 | 1040 | 0.00048295 | 0.99952
[6 empty] | 36 | 1050 | 0.00080943 | 0.99919
[6 empty] | 37 | 1060 | 0.0012383 | 0.99876
[6 empty] | 38 | 1070 | 0.024837 | 0.97516
[6 empty] | 39 | 1080 | 0.0053969 | 0.9946
[6 empty] | 40 | 1090 | 0.0063335 | 0.99367
[6 empty] | 41 | 1100 | 0.0029399 | 0.99706
[6 empty] | 42 | 1110 | 0.0021462 | 0.99785
[6 empty] | 43 | 1120 | 0.0028119 | 0.99719
[6 empty] | 44 | 1130 | 0.025735 | 0.97426
[6 empty] | 45 | 1140 | 0.0030973 | 0.9969
[6 empty] | 46 | 1150 | 0.0019923 | 0.99801
[6 empty] | 47 | 1160 | 0.0027296 | 0.99727
[6 empty] | 48 | 1170 | 0.0036926 | 0.99631
[6 empty] | 49 | 1180 | 0.0086683 | 0.99133
[6 empty] | 50 | 1190 | 0.0074988 | 0.9925
[6 empty] | 51 | 1200 | 0.0068277 | 0.99317
[6 empty] | 52 | 1210 | 0.014587 | 0.98541
[6 empty] | 53 | 1220 | 0.011504 | 0.9885
[6 empty] | 54 | 1230 | 0.0091951 | 0.9908
[6 empty] | 55 | 1240 | 0.0094014 | 0.9906
[6 empty] | 56 | 1250 | 0.0090129 | 0.99099
[6 empty] | 57 | 1260 | 0.009346 | 0.99065
[6 empty] | 58 | 1270 | 0.0071398 | 0.99286
[6 empty] | 59 | 1280 | 0.0061662 | 0.99383
[6 empty] | 60 | 1290 | 0.0078822 | 0.99212
[6 empty] | 61 | 1300 | 0.0079346 | 0.99207
[6 empty] | 62 | 1310 | 0.0057413 | 0.99426
[6 empty] | 63 | 1320 | 0.0057569 | 0.99424
[6 empty] | 64 | 1330 | 0.0090821 | 0.99092
[6 empty] | 65 | 1340 | 0.016208 | 0.98379
[6 empty] | 66 | 1350 | 0.015237 | 0.98476
[6 empty] | 67 | 1360 | 0.012031 | 0.98797
[6 empty] | 68 | 1370 | 0.010987 | 0.98901
[6 empty] | 69 | 1380 | 0.007489 | 0.99251
[6 empty] | 70 | 1390 | 0.0049048 | 0.9951
[6 empty] | 71 | 1400 | 0.0050855 | 0.99491
[6 empty] | 72 | 1410 | 0.010252 | 0.98975
[6 empty] | 73 | 1420 | 0.018778 | 0.98122
[6 empty] | 74 | 1430 | 0.0078974 | 0.9921
[6 empty] | 75 | 1440 | 0.0053934 | 0.99461
[6 empty] | 76 | 1450 | 0.0076593 | 0.99234
[6 empty] | 77 | 1460 | 0.017096 | 0.9829
[6 empty] | 78 | 1470 | 0.013837 | 0.98616
[6 empty] | 79 | 1480 | 0.0064591 | 0.99354
[6 empty] | 80 | 1490 | 0.0031618 | 0.99684
[6 empty] | 81 | 1500 | 0.001542 | 0.99846
[6 empty] | 82 | 1510 | 0.0009229 | 0.99908
[6 empty] | 83 | 1520 | 0.00078034 | 0.99922
[6 empty] | 84 | 1530 | 0.0010149 | 0.99899
[6 empty] | 85 | 1540 | 0.0024401 | 0.99756
[6 empty] | 86 | 1550 | 0.013815 | 0.98619
[6 empty] | 87 | 1560 | 0.0055091 | 0.99449
[6 empty] | 88 | 1570 | 0.0020895 | 0.99791
[6 empty] | 89 | 1580 | 0.0018744 | 0.99813
[6 empty] | 90 | 1590 | 0.0034163 | 0.99658
[6 empty] | 91 | 1600 | 0.0098182 | 0.99018
[6 empty] | 92 | 1610 | 0.0040265 | 0.99597
[6 empty] | 93 | 1620 | 0.001523 | 0.99848
[6 empty] | 94 | 1630 | 0.001048 | 0.99895
[6 empty] | 95 | 1640 | 0.001131 | 0.99887
[6 empty] | 96 | 1650 | 0.0018754 | 0.99812
[6 empty] | 97 | 1660 | 0.0050582 | 0.99494
[6 empty] | 98 | 1670 | 0.011942 | 0.98806
[6 empty] | 99 | 1680 | 0.0064091 | 0.99359
[6 empty] | 100 | 1690 | 0.0040472 | 0.99595
[6 empty] | 101 | 1700 | 0.0043033 | 0.9957
[6 empty] | 102 | 1710 | 0.0073632 | 0.99264
[6 empty] | 103 | 1720 | 0.014652 | 0.98535
[6 empty] | 104 | 1730 | 0.0096612 | 0.99034
[6 empty] | 105 | 1740 | 0.0045045 | 0.9955
[6 empty] | 106 | 1750 | 0.0029739 | 0.99703
[6 empty] | 107 | 1760 | 0.0026993 | 0.9973
[6 empty] | 108 | 1770 | 0.0032114 | 0.99679
[6 empty] | 109 | 1780 | 0.004961 | 0.99504
[6 empty] | 110 | 1790 | 0.010253 | 0.98975
[6 empty] | 111 | 1800 | 0.030459 | 0.96954

## OCAM
lambda [nm] | Lenslet | QE |  |  | lambda [nm] | QE (RAW)
700 | 0.99349 | 0.93603 |  |  | 700 | 0.938
710 | 0.99322 | 0.93877 |  |  | 750 | 0.932
720 | 0.99298 | 0.94016 |  |  | 800 | 0.904
730 | 0.99278 | 0.94021 |  |  | 850 | 0.839
740 | 0.99264 | 0.93892 |  |  | 900 | 0.734
750 | 0.99256 | 0.93629
760 | 0.99256 | 0.93231
770 | 0.99261 | 0.92699
780 | 0.9927 | 0.92033
790 | 0.99282 | 0.91232
800 | 0.99292 | 0.90297
810 | 0.99299 | 0.89228
820 | 0.99302 | 0.88025
830 | 0.993 | 0.86687
840 | 0.99291 | 0.85215
850 | 0.99274 | 0.83609
860 | 0.99248 | 0.81868
870 | 0.99212 | 0.79993
880 | 0.99164 | 0.77984
890 | 0.99104 | 0.75841
900 | 0.99033 | 0.73563
910 | 0.98949 | 0.71151
920 | 0.98851 | 0.68605
930 | 0.9874 | 0.65924 | AVG QE | 0.83596
940 | 0.98615 | 0.63109
950 | 0.98477 | 0.6016
960 | 0.98329 | 0.57077
970 | 0.98172 | 0.53859
980 | 0.98009 | 0.50507

## FL HA
Fiber throughput High Accuracy fiber
Fiber length [m] |  | 75 |  | AR coated surfaces: |  | 13
Stretcher losses |  | 0.9 |  | Glass air surface: |  | 1
FRD |  | 0.8 |  | Mirror surface: |  | 0
Efficiency correction |  | 4 | Values here gives a minimum absorption of 1dB/km (Thorlabs source);CERAMOPTEC the value is 4 dB/km for OH content = 1ppm; offeres 0.25ppm and 0.1ppm.
| Absorption in fiber |  | Fiber link Optics |  |  |  | measured
lbd[nm] | ABS [dB/km] | T | ARC 01 [R %] | ARC 02 [R %] | T optics calc. | T optics JL | Total HA
980 | 1.2263 | 0.86749 | 0.11906 | 0.16846 | 0.95214 | 0.83724 | 0.5986 | 0.7483
990 | 1.1876 | 0.8685 | 0.1192 | 0.13642 | 0.95212 | 0.84015 | 0.5994 | 0.74877
1000 | 1.1672 | 0.86903 | 0.13677 | 0.12533 | 0.94995 | 0.84261 | 0.6001 | 0.74924
1010 | 1.1542 | 0.86937 | 0.16676 | 0.12065 | 0.94625 | 0.84466 | 0.6008 | 0.74971
1020 | 1.1452 | 0.8696 | 0.20484 | 0.12771 | 0.94156 | 0.84632 | 0.6013 | 0.75018
1030 | 1.1385 | 0.86978 | 0.24871 | 0.14368 | 0.9362 | 0.8476 | 0.6018 | 0.75065
1040 | 1.1334 | 0.86991 | 0.28694 | 0.16315 | 0.93154 | 0.84852 | 0.6023 | 0.75112
1050 | 1.1295 | 0.87001 | 0.31077 | 0.18111 | 0.92865 | 0.84911 | 0.6026 | 0.75158
1060 | 1.1262 | 0.8701 | 0.32394 | 0.19424 | 0.92706 | 0.84939 | 0.6029 | 0.75205
1070 | 1.1232 | 0.87018 | 0.32812 | 0.20152 | 0.92655 | 0.84938 | 0.6032 | 0.75252
1080 | 1.1204 | 0.87025 | 0.32356 | 0.20455 | 0.92711 | 0.84909 | 0.6033 | 0.75299
1090 | 1.1173 | 0.87033 | 0.31124 | 0.20473 | 0.9286 | 0.84855 | 0.6035 | 0.75346
1100 | 1.1133 | 0.87044 | 0.29315 | 0.20296 | 0.93079 | 0.84779 | 0.6035 | 0.75393
1110 | 1.1078 | 0.87058 | 0.27108 | 0.19947 | 0.93347 | 0.84681 | 0.6035 | 0.7544
1120 | 1.099 | 0.87081 | 0.24656 | 0.19413 | 0.93646 | 0.84565 | 0.6035 | 0.75355
1130 | 1.0834 | 0.87122 | 0.22076 | 0.18691 | 0.93961 | 0.84432 | 0.6034 | 0.7527
1140 | 1.0544 | 0.87198 | 0.19455 | 0.17792 | 0.94283 | 0.84284 | 0.6032 | 0.75185
1150 | 1.0073 | 0.87321 | 0.16864 | 0.1679 | 0.94601 | 0.84123 | 0.603 | 0.75101
1160 | 0.9774 | 0.87399 | 0.14375 | 0.15782 | 0.94908 | 0.83952 | 0.6027 | 0.75016
1170 | 0.9663 | 0.87428 | 0.12063 | 0.14861 | 0.95194 | 0.83773 | 0.6024 | 0.74931
1180 | 0.9623 | 0.87439 | 0.10027 | 0.14145 | 0.95447 | 0.83587 | 0.6021 | 0.74846
1190 | 0.9624 | 0.87439 | 0.08376 | 0.13763 | 0.95652 | 0.83396 | 0.6017 | 0.74761
1200 | 0.9656 | 0.8743 | 0.07209 | 0.13777 | 0.95798 | 0.83204 | 0.6013 | 0.74676
1210 | 0.9715 | 0.87415 | 0.06602 | 0.14175 | 0.95873 | 0.83011 | 0.6008 | 0.74591
1220 | 0.9802 | 0.87392 | 0.06591 | 0.14884 | 0.95875 | 0.8282 | 0.6003 | 0.74506
1230 | 0.9924 | 0.8736 | 0.07095 | 0.15823 | 0.95812 | 0.82633 | 0.5997 | 0.74422
1240 | 1.0097 | 0.87315 | 0.08026 | 0.16943 | 0.95696 | 0.82453 | 0.5991 | 0.74337
1250 | 1.0365 | 0.87245 | 0.09338 | 0.18209 | 0.95533 | 0.8228 | 0.5985 | 0.74252
1260 | 1.0867 | 0.87113 | 0.11031 | 0.19594 | 0.95322 | 0.82117 | 0.5979 | 0.74167
1270 | 1.3767 | 0.86359 | 0.13202 | 0.21077 | 0.95053 | 0.81964 | 0.5972 | 0.74082
1280 | 1.4034 | 0.86289 | 0.15789 | 0.22638 | 0.94734 | 0.81822 | 0.5965 | 0.73997
1290 | 1.3615 | 0.86398 | 0.18514 | 0.24255 | 0.94398 | 0.8169 | 0.5957 | 0.73912
1300 | 1.2951 | 0.8657 | 0.21248 | 0.25913 | 0.94063 | 0.81569 | 0.595 | 0.73827
1310 | 1.1905 | 0.86842 | 0.2396 | 0.27603 | 0.93731 | 0.81459 | 0.5942 | 0.73743
1320 | 1.063 | 0.87175 | 0.2665 | 0.29312 | 0.93403 | 0.81359 | 0.5934 | 0.73658
1330 | 1.0563 | 0.87193 | 0.29321 | 0.31022 | 0.93078 | 0.8127 | 0.5925 | 0.73573
1340 | 1.1148 | 0.8704 | 0.31974 | 0.32652 | 0.92757 | 0.81193 | 0.5917 | 0.73488
1350 | 1.2102 | 0.86791 | 0.34609 | 0.3408 | 0.92439 | 0.81126 | 0.5908 | 0.73403
1360 | 1.3401 | 0.86453 | 0.37223 | 0.35235 | 0.92124 | 0.81071 | 0.5899 | 0.73318
1370 | 1.5114 | 0.8601 | 0.39814 | 0.36127 | 0.91813 | 0.81027 | 0.589 | 0.73233
1380 | 1.7601 | 0.85371 | 0.42373 | 0.36804 | 0.91507 | 0.80994 | 0.5881 | 0.73148
1390 | 2.1798 | 0.84303 | 0.4489 | 0.37295 | 0.91206 | 0.80973 | 0.5872 | 0.73064
1400 | 3.241 | 0.81661 | 0.47337 | 0.37606 | 0.90915 | 0.80964 | 0.5863 | 0.72979
1410 | 8.289 | 0.70185 | 0.4964 | 0.3773 | 0.90642 | 0.80966 | 0.5853 | 0.72894
1420 | 9.586 | 0.67507 | 0.5155 | 0.37652 | 0.90416 | 0.8098 | 0.5844 | 0.72809
1430 | 7.46 | 0.71953 | 0.52104 | 0.37348 | 0.90351 | 0.81007 | 0.5834 | 0.72724
1440 | 4.482 | 0.78677 | 0.51372 | 0.36781 | 0.90437 | 0.81045 | 0.5825 | 0.72639
1450 | 2.701 | 0.82995 | 0.50732 | 0.35905 | 0.90513 | 0.81095 | 0.5815 | 0.72554
1460 | 1.9024 | 0.85007 | 0.49998 | 0.34727 | 0.906 | 0.81158 | 0.5806 | 0.72469
1470 | 1.5374 | 0.85943 | 0.49083 | 0.33293 | 0.90708 | 0.81233 | 0.5796 | 0.72385
1480 | 1.3071 | 0.86539 | 0.47925 | 0.31655 | 0.90846 | 0.81321 | 0.5787 | 0.723
1490 | 1.1396 | 0.86975 | 0.46494 | 0.29879 | 0.91016 | 0.81421 | 0.5777 | 0.72215
1500 | 1.0181 | 0.87293 | 0.44832 | 0.28061 | 0.91213 | 0.81534 | 0.5768 | 0.7213 | 1.2505
1510 | 0.9275 | 0.8753 | 0.42987 | 0.26344 | 0.91433 | 0.81659 | 0.5759 | 0.72017
1520 | 0.8587 | 0.87711 | 0.41005 | 0.24832 | 0.9167 | 0.81795 | 0.575 | 0.71905
1530 | 0.8087 | 0.87843 | 0.38923 | 0.23466 | 0.9192 | 0.81939 | 0.5741 | 0.71792
1540 | 0.7732 | 0.87936 | 0.36771 | 0.22109 | 0.92178 | 0.82088 | 0.5732 | 0.7168
1550 | 0.7485 | 0.88002 | 0.34573 | 0.2062 | 0.92443 | 0.82239 | 0.5723 | 0.71567
1560 | 0.7318 | 0.88046 | 0.3235 | 0.18943 | 0.92711 | 0.82392 | 0.5714 | 0.71455
1570 | 0.7211 | 0.88074 | 0.30116 | 0.17188 | 0.92982 | 0.82542 | 0.5706 | 0.71355
1580 | 0.715 | 0.8809 | 0.27883 | 0.15469 | 0.93253 | 0.82687 | 0.5698 | 0.71255
1590 | 0.7127 | 0.88096 | 0.25661 | 0.13872 | 0.93523 | 0.82826 | 0.569 | 0.71155
1600 | 0.7134 | 0.88094 | 0.23461 | 0.12478 | 0.93792 | 0.82955 | 0.5682 | 0.71055
1610 | 0.7164 | 0.88086 | 0.21291 | 0.11372 | 0.94057 | 0.83073 | 0.5674 | 0.70955
1620 | 0.7214 | 0.88073 | 0.19162 | 0.10616 | 0.94319 | 0.83175 | 0.5667 | 0.70867
1630 | 0.7282 | 0.88055 | 0.17084 | 0.10178 | 0.94574 | 0.83261 | 0.566 | 0.70779
1640 | 0.7366 | 0.88033 | 0.15092 | 0.1002 | 0.9482 | 0.83328 | 0.5653 | 0.70692
1650 | 0.7466 | 0.88007 | 0.13233 | 0.10104 | 0.9505 | 0.83373 | 0.5647 | 0.70617
1660 | 0.7581 | 0.87976 | 0.1156 | 0.10391 | 0.95257 | 0.83393 | 0.5641 | 0.70542
1670 | 0.7712 | 0.87942 | 0.10129 | 0.10866 | 0.95434 | 0.83387 | 0.5635 | 0.70467
1680 | 0.7859 | 0.87903 | 0.09002 | 0.11784 | 0.95574 | 0.83351 | 0.563 | 0.70404
1690 | 0.8025 | 0.87859 | 0.08246 | 0.13397 | 0.95669 | 0.83283 | 0.5625 | 0.70342
1700 | 0.821 | 0.8781 | 0.07932 | 0.1567 | 0.95708 | 0.83182 | 0.562 | 0.70279
1710 | 0.8418 | 0.87756 | 0.08142 | 0.18422 | 0.95681 | 0.83043 | 0.5616 | 0.70229
1720 | 0.8654 | 0.87693 | 0.08936 | 0.21557 | 0.95583 | 0.82866 | 0.5613 | 0.70192
1730 | 0.8926 | 0.87622 | 0.10325 | 0.25072 | 0.9541 | 0.82647 | 0.5609 | 0.70142
1740 | 0.925 | 0.87537 | 0.12302 | 0.28974 | 0.95165 | 0.82383 | 0.5606 | 0.70104
1750 | 0.9655 | 0.87431 | 0.14853 | 0.33275 | 0.94849 | 0.82073 | 0.5604 | 0.70079
1760 | 1.019 | 0.8729 | 0.17955 | 0.37994 | 0.94467 | 0.81714 | 0.5602 | 0.70054
1770 | 1.0942 | 0.87094 | 0.21576 | 0.43157 | 0.94023 | 0.81304 | 0.5601 | 0.70042
1780 | 1.2033 | 0.86809 | 0.25677 | 0.48793 | 0.93521 | 0.80839 | 0.56 | 0.70029
1790 | 1.3546 | 0.86416 | 0.30214 | 0.54919 | 0.9297 | 0.80318 | 0.56 | 0.70029
1800 | 1.5381 | 0.85941 | 0.35142 | 0.61506 | 0.92374 | 0.79738 | 0.56 | 0.70029

## FL HE
Fiber throughput High Efficiency fiber
Fiber length [m] |  | 45 |  | AR coated surfaces |  | 13
FRD |  | 0.8 |  | Glass-air surfaces |  | 1
Efficiency correction |  | 4 |  | Mirror surfaces |  | 2
Values here gives a minimum absorption of 1dB/km (Thorlabs source);CERAMOPTEC the value is 4 dB/km for OH content = 1ppm; offeres 0.25ppm and 0.1ppm.
| Absorption in fiber |  | Fiber link Optics |  |  |  | measured
lbd[nm] | ABS [dB/km] | T | ARC 01 [R %] | ARC 02 [R %] | T optics cal. | T optics JL | Total HE
980 | 1.2263 | 0.97817 | 0.11906 | 0.16846 | 0.9106 | 0.80635 | 0.572 | 0.715
990 | 1.1876 | 0.97885 | 0.1192 | 0.13642 | 0.91449 | 0.80875 | 0.573 | 0.71555
1000 | 1.1672 | 0.97921 | 0.13677 | 0.12533 | 0.91589 | 0.81077 | 0.5738 | 0.71611
1010 | 1.1542 | 0.97944 | 0.16676 | 0.12065 | 0.91675 | 0.81241 | 0.5746 | 0.71666
1020 | 1.1452 | 0.9796 | 0.20484 | 0.12771 | 0.91614 | 0.81372 | 0.5753 | 0.71722
1030 | 1.1385 | 0.97972 | 0.24871 | 0.14368 | 0.9142 | 0.81469 | 0.5759 | 0.71777
1040 | 1.1334 | 0.97981 | 0.28694 | 0.16315 | 0.9113 | 0.81535 | 0.5764 | 0.71832
1050 | 1.1295 | 0.97987 | 0.31077 | 0.18111 | 0.9099 | 0.81572 | 0.5768 | 0.71888
1060 | 1.1262 | 0.97993 | 0.32394 | 0.19424 | 0.90858 | 0.81581 | 0.5772 | 0.71943
1070 | 1.1232 | 0.97999 | 0.32812 | 0.20152 | 0.90744 | 0.81565 | 0.5774 | 0.71998
1080 | 1.1204 | 0.98003 | 0.32356 | 0.20455 | 0.90687 | 0.81525 | 0.5776 | 0.72054
1090 | 1.1173 | 0.98009 | 0.31124 | 0.20473 | 0.90704 | 0.81463 | 0.5777 | 0.72109
1100 | 1.1133 | 0.98016 | 0.29315 | 0.20296 | 0.90672 | 0.8138 | 0.5778 | 0.72165
1110 | 1.1078 | 0.98026 | 0.27108 | 0.19947 | 0.90721 | 0.8128 | 0.5777 | 0.7222
1120 | 1.099 | 0.98041 | 0.24656 | 0.19413 | 0.90764 | 0.81163 | 0.5776 | 0.72112
1130 | 1.0834 | 0.98069 | 0.22076 | 0.18691 | 0.90869 | 0.81032 | 0.5774 | 0.72004
1140 | 1.0544 | 0.9812 | 0.19455 | 0.17792 | 0.91023 | 0.80888 | 0.5772 | 0.71895
1150 | 1.0073 | 0.98203 | 0.16864 | 0.1679 | 0.91165 | 0.80732 | 0.5769 | 0.71787
1160 | 0.9774 | 0.98256 | 0.14375 | 0.15782 | 0.91239 | 0.80568 | 0.5766 | 0.71679
1170 | 0.9663 | 0.98276 | 0.12063 | 0.14861 | 0.91296 | 0.80397 | 0.5761 | 0.71571
1180 | 0.9623 | 0.98283 | 0.10027 | 0.14145 | 0.91392 | 0.8022 | 0.5757 | 0.71463
1190 | 0.9624 | 0.98283 | 0.08376 | 0.13763 | 0.91471 | 0.80039 | 0.5751 | 0.71354
1200 | 0.9656 | 0.98277 | 0.07209 | 0.13777 | 0.91544 | 0.79856 | 0.5746 | 0.71246
1210 | 0.9715 | 0.98267 | 0.06602 | 0.14175 | 0.91359 | 0.79674 | 0.5739 | 0.71138
1220 | 0.9802 | 0.98251 | 0.06591 | 0.14884 | 0.91259 | 0.79494 | 0.5732 | 0.7103
1230 | 0.9924 | 0.9823 | 0.07095 | 0.15823 | 0.91119 | 0.79317 | 0.5725 | 0.70922
1240 | 1.0097 | 0.98199 | 0.08026 | 0.16943 | 0.91076 | 0.79145 | 0.5718 | 0.70813
1250 | 1.0365 | 0.98152 | 0.09338 | 0.18209 | 0.90946 | 0.78982 | 0.571 | 0.70705
1260 | 1.0867 | 0.98063 | 0.11031 | 0.19594 | 0.90836 | 0.78827 | 0.5701 | 0.70597
1270 | 1.3767 | 0.97552 | 0.13202 | 0.21077 | 0.90626 | 0.78681 | 0.5692 | 0.70489
1280 | 1.4034 | 0.97506 | 0.15789 | 0.22638 | 0.90489 | 0.78545 | 0.5683 | 0.70381
1290 | 1.3615 | 0.97579 | 0.18514 | 0.24255 | 0.9027 | 0.78418 | 0.5674 | 0.70272
1300 | 1.2951 | 0.97696 | 0.21248 | 0.25913 | 0.90037 | 0.78302 | 0.5664 | 0.70164
1310 | 1.1905 | 0.9788 | 0.2396 | 0.27603 | 0.89897 | 0.78195 | 0.5654 | 0.70056
1320 | 1.063 | 0.98105 | 0.2665 | 0.29312 | 0.89645 | 0.78098 | 0.5644 | 0.69948
1330 | 1.0563 | 0.98117 | 0.29321 | 0.31022 | 0.89452 | 0.78012 | 0.5633 | 0.69839
1340 | 1.1148 | 0.98013 | 0.31974 | 0.32652 | 0.89207 | 0.77935 | 0.5622 | 0.69731
1350 | 1.2102 | 0.97845 | 0.34609 | 0.3408 | 0.89 | 0.7787 | 0.5611 | 0.69623
1360 | 1.3401 | 0.97617 | 0.37223 | 0.35235 | 0.88932 | 0.77814 | 0.56 | 0.69515
1370 | 1.5114 | 0.97316 | 0.39814 | 0.36127 | 0.88972 | 0.7777 | 0.5589 | 0.69407
1380 | 1.7601 | 0.96881 | 0.42373 | 0.36804 | 0.88847 | 0.77737 | 0.5578 | 0.69298
1390 | 2.1798 | 0.96152 | 0.4489 | 0.37295 | 0.8881 | 0.77714 | 0.5566 | 0.6919
1400 | 3.241 | 0.94333 | 0.47337 | 0.37606 | 0.8878 | 0.77703 | 0.5555 | 0.69082
1410 | 8.289 | 0.8614 | 0.4964 | 0.3773 | 0.88815 | 0.77703 | 0.5543 | 0.68974
1420 | 9.586 | 0.84152 | 0.5155 | 0.37652 | 0.88796 | 0.77714 | 0.5532 | 0.68866
1430 | 7.46 | 0.87435 | 0.52104 | 0.37348 | 0.88778 | 0.77737 | 0.552 | 0.68757
1440 | 4.482 | 0.92249 | 0.51372 | 0.36781 | 0.8887 | 0.77772 | 0.5508 | 0.68649
1450 | 2.701 | 0.95254 | 0.50732 | 0.35905 | 0.89003 | 0.77819 | 0.5497 | 0.68541
1460 | 1.9024 | 0.96634 | 0.49998 | 0.34727 | 0.89095 | 0.77877 | 0.5485 | 0.68433
1470 | 1.5374 | 0.97271 | 0.49083 | 0.33293 | 0.89349 | 0.77948 | 0.5474 | 0.68325
1480 | 1.3071 | 0.97675 | 0.47925 | 0.31655 | 0.89512 | 0.78031 | 0.5462 | 0.68216
1490 | 1.1396 | 0.9797 | 0.46494 | 0.29879 | 0.89675 | 0.78126 | 0.5451 | 0.68108
1500 | 1.0181 | 0.98184 | 0.44832 | 0.28061 | 0.89879 | 0.78235 | 0.544 | 0.68 | 1.25
1510 | 0.9275 | 0.98344 | 0.42987 | 0.26344 | 0.90047 | 0.78355 | 0.5429 | 0.67863
1520 | 0.8587 | 0.98466 | 0.41005 | 0.24832 | 0.90146 | 0.78485 | 0.5418 | 0.67725
1530 | 0.8087 | 0.98555 | 0.38923 | 0.23466 | 0.90287 | 0.78624 | 0.5408 | 0.676
1540 | 0.7732 | 0.98618 | 0.36771 | 0.22109 | 0.90453 | 0.78768 | 0.5397 | 0.67462
1550 | 0.7485 | 0.98662 | 0.34573 | 0.2062 | 0.90674 | 0.78915 | 0.5387 | 0.67337
1560 | 0.7318 | 0.98691 | 0.3235 | 0.18943 | 0.90832 | 0.79062 | 0.5377 | 0.67212
1570 | 0.7211 | 0.9871 | 0.30116 | 0.17188 | 0.91018 | 0.79208 | 0.5368 | 0.671
1580 | 0.715 | 0.98721 | 0.27883 | 0.15469 | 0.91229 | 0.7935 | 0.5359 | 0.66988
1590 | 0.7127 | 0.98725 | 0.25661 | 0.13872 | 0.91414 | 0.79485 | 0.535 | 0.66875
1600 | 0.7134 | 0.98724 | 0.23461 | 0.12478 | 0.91577 | 0.79611 | 0.5341 | 0.66763
1610 | 0.7164 | 0.98719 | 0.21291 | 0.11372 | 0.91786 | 0.79725 | 0.5333 | 0.66663
1620 | 0.7214 | 0.9871 | 0.19162 | 0.10616 | 0.91913 | 0.79826 | 0.5325 | 0.66563
1630 | 0.7282 | 0.98698 | 0.17084 | 0.10178 | 0.92017 | 0.7991 | 0.5318 | 0.66475
1640 | 0.7366 | 0.98683 | 0.15092 | 0.1002 | 0.92069 | 0.79976 | 0.5311 | 0.66387
1650 | 0.7466 | 0.98665 | 0.13233 | 0.10104 | 0.92104 | 0.8002 | 0.5305 | 0.66312
1660 | 0.7581 | 0.98645 | 0.1156 | 0.10391 | 0.92118 | 0.80041 | 0.5299 | 0.66238
1670 | 0.7712 | 0.98621 | 0.10129 | 0.10866 | 0.9209 | 0.80036 | 0.5294 | 0.66175
1680 | 0.7859 | 0.98595 | 0.09002 | 0.11784 | 0.91863 | 0.80002 | 0.5289 | 0.66112
1690 | 0.8025 | 0.98566 | 0.08246 | 0.13397 | 0.91652 | 0.79938 | 0.5284 | 0.6605
1700 | 0.821 | 0.98533 | 0.07932 | 0.1567 | 0.9139 | 0.79841 | 0.5281 | 0.66013
1710 | 0.8418 | 0.98496 | 0.08142 | 0.18422 | 0.91063 | 0.79708 | 0.5278 | 0.65975
1720 | 0.8654 | 0.98454 | 0.08936 | 0.21557 | 0.90598 | 0.79536 | 0.5275 | 0.65937
1730 | 0.8926 | 0.98406 | 0.10325 | 0.25072 | 0.90194 | 0.79325 | 0.5273 | 0.65912
1740 | 0.925 | 0.98349 | 0.12302 | 0.28974 | 0.89701 | 0.7907 | 0.5272 | 0.659
1750 | 0.9655 | 0.98277 | 0.14853 | 0.33275 | 0.89222 | 0.7877 | 0.5271 | 0.65887
1760 | 1.019 | 0.98183 | 0.17955 | 0.37994 | 0.88627 | 0.78423 | 0.5272 | 0.659
1770 | 1.0942 | 0.9805 | 0.21576 | 0.43157 | 0.87965 | 0.78025 | 0.5273 | 0.65912
1780 | 1.2033 | 0.97857 | 0.25677 | 0.48793 | 0.87373 | 0.77575 | 0.5274 | 0.65925
1790 | 1.3546 | 0.97591 | 0.30214 | 0.54919 | 0.86682 | 0.7707 | 0.5277 | 0.65962
1800 | 1.5381 | 0.97269 | 0.35142 | 0.61506 | 0.85861 | 0.76507 | 0.528 | 0.66

## BE
YJH spectrometer
Data from "NIRPS BE transmission As-built.xls" (June 2021)
lbd [nm] | IJH spectrograph | Column1 | Column2
980 | 0.17403 |  | 17.403
990 | 0.17593 |  | 17.5934
1000 | 0.1781 |  | 17.8101
1010 | 0.18135 |  | 18.1354
1020 | 0.1854 |  | 18.5401
1030 | 0.19022 |  | 19.0225
1040 | 0.19599 |  | 19.599
1050 | 0.20225 |  | 20.2252
1060 | 0.20871 |  | 20.8709
1070 | 0.21527 |  | 21.5275
1080 | 0.22116 |  | 22.1163
1090 | 0.22614 |  | 22.6138
1100 | 0.23084 |  | 23.0837
1110 | 0.23583 |  | 23.5827
1120 | 0.23958 |  | 23.9584
1130 | 0.24264 |  | 24.2638
1140 | 0.24491 |  | 24.4907
1150 | 0.24669 |  | 24.6685
1160 | 0.24854 |  | 24.8542
1170 | 0.2504 |  | 25.0403
1180 | 0.25217 |  | 25.2172
1190 | 0.25303 |  | 25.3031
1200 | 0.25354 |  | 25.3537
1210 | 0.25503 |  | 25.503
1220 | 0.25694 |  | 25.6945
1230 | 0.25839 |  | 25.8387
1240 | 0.25847 |  | 25.8466
1250 | 0.25906 |  | 25.9061
1260 | 0.26075 |  | 26.0752
1270 | 0.26284 |  | 26.2843
1280 | 0.26505 |  | 26.5046
1290 | 0.26713 |  | 26.7132
1300 | 0.26939 |  | 26.9387
1310 | 0.27086 |  | 27.0863
1320 | 0.27183 |  | 27.1833
1330 | 0.27348 |  | 27.3482
1340 | 0.27599 |  | 27.5987
1350 | 0.27899 |  | 27.8993
1360 | 0.27988 |  | 27.9884
1370 | 0.27894 |  | 27.8936
1380 | 0.27719 |  | 27.7187
1390 | 0.27601 |  | 27.6015
1400 | 0.2756 |  | 27.5604
1410 | 0.27569 |  | 27.5685
1420 | 0.27631 |  | 27.6307
1430 | 0.27667 |  | 27.6672
1440 | 0.27773 |  | 27.7734
1450 | 0.27895 |  | 27.8953
1460 | 0.28029 |  | 28.0295
1470 | 0.28123 |  | 28.1226
1480 | 0.28076 |  | 28.0761
1490 | 0.27858 |  | 27.8575
1500 | 0.27693 |  | 27.6931
1510 | 0.27758 |  | 27.7582
1520 | 0.27907 |  | 27.9069
1530 | 0.28206 |  | 28.2058
1540 | 0.28386 |  | 28.3857
1550 | 0.28601 |  | 28.6011
1560 | 0.28751 |  | 28.7508
1570 | 0.28931 |  | 28.9313
1580 | 0.29192 |  | 29.1925
1590 | 0.29472 |  | 29.472
1600 | 0.29695 |  | 29.6951
1610 | 0.29893 |  | 29.8931
1620 | 0.30055 |  | 30.0546
1630 | 0.30197 |  | 30.1974
1640 | 0.30413 |  | 30.4125
1650 | 0.30613 |  | 30.613
1660 | 0.30883 |  | 30.8833
1670 | 0.31084 |  | 31.0837
1680 | 0.31247 |  | 31.2466
1690 | 0.3136 |  | 31.3603
1700 | 0.31398 |  | 31.3981
1710 | 0.31375 |  | 31.3749
1720 | 0.31321 |  | 31.3205
1730 | 0.31246 |  | 31.2457
1740 | 0.31137 |  | 31.137
1750 | 0.30998 |  | 30.9979
1760 | 0.30836 |  | 30.8361
1770 | 0.30571 |  | 30.5711
1780 | 0.3024 |  | 30.2395
1790 | 0.29843 |  | 29.8433
1800 | 0.29365 |  | 29.3646

## H4RG
H4RG QE
RAW from E. Artigau mail on 01.06.2017
lbd [nm] | H4RG [RAW]
800 | 93
1000 | 96
1230 | 98
1500 | 99
2000 | 79
Interpolated values from 3rd order polynomial fit
lbd [nm] | H4RG
980 | 0.92218
990 | 0.92304
1000 | 0.9239
1010 | 0.92476
1020 | 0.92562
1030 | 0.92647
1040 | 0.92732
1050 | 0.92816
1060 | 0.929
1070 | 0.92982
1080 | 0.93063
1090 | 0.93142
1100 | 0.9322
1110 | 0.93296
1120 | 0.9337
1130 | 0.93442
1140 | 0.93512
1150 | 0.93579
1160 | 0.93643
1170 | 0.93705
1180 | 0.93763
1190 | 0.93818
1200 | 0.9387
1210 | 0.93918
1220 | 0.93963
1230 | 0.94003
1240 | 0.94039
1250 | 0.94071
1260 | 0.94099
1270 | 0.94122
1280 | 0.94139
1290 | 0.94152
1300 | 0.9416
1310 | 0.94162
1320 | 0.94159
1330 | 0.9415
1340 | 0.94135
1350 | 0.94114
1360 | 0.94086
1370 | 0.94052
1380 | 0.94012
1390 | 0.93964
1400 | 0.9391
1410 | 0.93848
1420 | 0.93779
1430 | 0.93703
1440 | 0.93618
1450 | 0.93526
1460 | 0.93426
1470 | 0.93317
1480 | 0.932
1490 | 0.93075
1500 | 0.9294
1510 | 0.92796
1520 | 0.92644
1530 | 0.92482
1540 | 0.9231
1550 | 0.92129
1560 | 0.91938
1570 | 0.91736
1580 | 0.91525
1590 | 0.91303
1600 | 0.9107
1610 | 0.90827
1620 | 0.90572
1630 | 0.90307
1640 | 0.9003
1650 | 0.89741
1660 | 0.89441
1670 | 0.89129
1680 | 0.88805
1690 | 0.88469
1700 | 0.8812
1710 | 0.87759
1720 | 0.87385
1730 | 0.86997
1740 | 0.86597
1750 | 0.86184
1760 | 0.85757
1770 | 0.85316
1780 | 0.84861
1790 | 0.84393
1800 | 0.8391

## DM241
LBD [nm] | R [%]
700 | 97.2879
710 | 97.2828
720 | 97.2875
730 | 97.2531
740 | 97.2739
750 | 97.31
760 | 97.2731
770 | 97.3259
780 | 97.3029
790 | 97.2465
800 | 97.2704
810 | 97.8163
820 | 97.4883
830 | 97.4595
840 | 97.5131
850 | 97.5823
860 | 97.7074
870 | 97.6282
880 | 97.6207
890 | 97.8969
900 | 97.8538
910 | 98.0756
920 | 98.0261
930 | 98.1109
940 | 98.0812
950 | 98.1211
960 | 98.1504
970 | 98.1697
980 | 98.1783
990 | 98.2586
1000 | 98.248
1010 | 98.2738
1020 | 98.376
1030 | 98.3078
1040 | 98.3489
1050 | 98.2379
1060 | 98.3254
1070 | 98.4668
1080 | 98.4277
1090 | 98.4361
1100 | 98.3848
1110 | 98.5678
1120 | 98.5202
1130 | 98.546
1140 | 98.5802
1150 | 98.5922
1160 | 98.6268
1170 | 98.6285
1180 | 98.551
1190 | 98.6057
1200 | 98.6768
1210 | 98.6982
1220 | 98.6546
1230 | 98.721
1240 | 98.6986
1250 | 98.7207
1260 | 98.6993
1270 | 98.751
1280 | 98.8209
1290 | 98.8444
1300 | 98.7396
1310 | 98.7993
1320 | 98.8013
1330 | 98.7387
1340 | 98.8374
1350 | 98.8025
1360 | 98.8435
1370 | 98.9202
1380 | 98.7821
1390 | 98.8704
1400 | 98.8496
1410 | 98.9482
1420 | 98.8768
1430 | 98.9419
1440 | 98.9066
1450 | 98.9097
1460 | 98.8942
1470 | 98.86
1480 | 98.9796
1490 | 98.9016
1500 | 98.9323
1510 | 98.9462
1520 | 98.9295
1530 | 99.0129
1540 | 99.0501
1550 | 98.9585
1560 | 99.0038
1570 | 99.0335
1580 | 99.0725
1590 | 99.0971
1600 | 99.0696
1610 | 99.1363
1620 | 99.0257
1630 | 99.0936
1640 | 99.0642
1650 | 99.0171
1660 | 99.0332
1670 | 99.0423
1680 | 99.0935
1690 | 99.0729
1700 | 99.1537
1710 | 99.1881
1720 | 99.1361
1730 | 99.1498
1740 | 99.1349
1750 | 99.1555
1760 | 98.9988
1770 | 99.1103
1780 | 99.189
1790 | 99.2056
1800 | 99.1989
1810 | 99.2017
1820 | 99.2747
1830 | 99.2008
1840 | 99.157
1850 | 99.1729
1860 | 99.1075
1870 | 99.0961
1880 | 99.1331
1890 | 99.0661
1900 | 99.1378
1910 | 99.1637
1920 | 99.187
1930 | 99.1429
1940 | 99.1628
1950 | 99.2118
1960 | 99.1759
1970 | 99.1592
1980 | 99.2448
1990 | 99.0961
2000 | 99.0886
2010 | 99.0665
2020 | 99.025
2030 | 99.0818
2040 | 99.0708
2050 | 99.1494
2060 | 99.1571
2070 | 99.2091
2080 | 99.2972
2090 | 99.2916
2100 | 99.2959
2110 | 99.3565
2120 | 99.3376
2130 | 99.268
2140 | 99.2032
2150 | 99.1142
2160 | 99.0745
2170 | 99.0319
2180 | 99.0989
2190 | 99.0071
2200 | 99.0171
2210 | 99.0446
2220 | 99.05
2230 | 99.1275
2240 | 99.2161
2250 | 99.1799
2260 | 99.3274
2270 | 99.3318
2280 | 99.2354
2290 | 99.2958
2300 | 99.2479
2310 | 99.2857
2320 | 99.2021
2330 | 99.2729
2340 | 99.188
2350 | 99.1221
2360 | 99.0302
2370 | 99.1731
2380 | 99.0984
2390 | 99.0506
2400 | 99.0408

## OPTIMAX coating
All AR coatings from OPTIMAX
+ We use only one curve of the coatings measurement provided by Optimax for WFS and YJH arms, as they are very close form each other. + The values reported here are the transmision per sruface.
lbd [nm] | ADC 5381 | ADC 5372 | ADC 5354 | BBAR WFS | BBAR YJH | BBAR YJH2 | Column3
700 | 0.99457 | 0.99622 | 0.99714 | 0.99879
710 | 0.99754 | 0.99625 | 0.99561 | 0.99951
720 | 0.9974 | 0.99519 | 0.99365 | 0.99962
730 | 0.99584 | 0.99391 | 0.99167 | 0.99956
740 | 0.99418 | 0.99288 | 0.99047 | 0.99947
750 | 0.99282 | 0.99238 | 0.99036 | 0.99931
760 | 0.99203 | 0.99224 | 0.99082 | 0.99905
770 | 0.992 | 0.99228 | 0.99155 | 0.99872
780 | 0.99252 | 0.99238 | 0.99228 | 0.99838
790 | 0.99324 | 0.99247 | 0.99293 | 0.9981
800 | 0.99397 | 0.99253 | 0.99343 | 0.9979
810 | 0.99462 | 0.99251 | 0.99365 | 0.99781
820 | 0.99516 | 0.99237 | 0.99356 | 0.99777
830 | 0.99551 | 0.99207 | 0.9932 | 0.99779
840 | 0.99554 | 0.99164 | 0.99267 | 0.99786
850 | 0.99522 | 0.99115 | 0.99205 | 0.99795
860 | 0.99466 | 0.99066 | 0.99142 | 0.99808
870 | 0.994 | 0.99024 | 0.99081 | 0.99822
880 | 0.99334 | 0.98992 | 0.99027 | 0.99839
890 | 0.99269 | 0.9897 | 0.98984 | 0.99856
900 | 0.99206 | 0.98958 | 0.98953 | 0.99873 | 0.97246
910 | 0.99147 | 0.98952 | 0.98938 | 0.99888 | 0.98537
920 | 0.99094 | 0.98951 | 0.9894 | 0.99899 | 0.99281
930 | 0.9905 | 0.98954 | 0.98954 | 0.99905 | 0.99731
940 | 0.99018 | 0.9896 | 0.98975 | 0.99903 | 0.99883
950 | 0.98998 | 0.98969 | 0.99 | 0.99892 | 0.99894
960 | 0.98989 | 0.98981 | 0.99024 | 0.99873 | 0.99829
970 | 0.9899 | 0.98995 | 0.99046 | 0.99846 | 0.99725
980 | 0.98998 | 0.99013 | 0.99066 | 0.99812 | 0.99628
990 | 0.9901 | 0.99035 | 0.99085 | 0.9977 | 0.99556
1000 | 0.99025 | 0.99049 | 0.991 | 0.9972 | 0.9951
1010 | 0.99043 | 0.99059 | 0.99114 | 0.99662 | 0.995
1020 | 0.99063 | 0.99066 | 0.99124 | 0.99593 | 0.99526
1030 | 0.99084 | 0.99072 | 0.99129 | 0.99513 | 0.99576
1040 | 0.99106 | 0.99075 | 0.99129 | 0.99422 | 0.99636
1050 | 0.99126 | 0.99076 | 0.99121 | 0.99319 | 0.99694
1060 | 0.99144 | 0.99073 | 0.99106 | 0.99203 | 0.99745
1070 | 0.99159 | 0.99067 | 0.99083 | 0.99076 | 0.99788
1080 | 0.99171 | 0.99056 | 0.99056 | 0.9894 | 0.99818
1090 | 0.99177 | 0.99042 | 0.99026 | 0.98798 | 0.99831
1100 | 0.9918 | 0.99026 | 0.98996 | 0.98655 | 0.99826
1110 | 0.99179 | 0.99007 | 0.98967 |  | 0.99806
1120 | 0.99175 | 0.98988 | 0.98937 |  | 0.99775
1130 | 0.99167 | 0.98967 | 0.98909 |  | 0.99739
1140 | 0.99157 | 0.98947 | 0.98881 |  | 0.99701
1150 | 0.99144 | 0.98926 | 0.98855 |  | 0.99663
1160 | 0.9913 | 0.98906 | 0.9883 |  | 0.99627
1170 | 0.99115 | 0.98887 | 0.98807 |  | 0.99598
1180 | 0.99097 | 0.98869 | 0.98786 |  | 0.99576
1190 | 0.99079 | 0.98852 | 0.98768 |  | 0.99565
1200 | 0.9906 | 0.98836 | 0.98753 |  | 0.99565
1210 | 0.9904 | 0.98821 | 0.98739 |  | 0.99572
1220 | 0.99019 | 0.98808 | 0.98728 |  | 0.99587
1230 | 0.98998 | 0.98795 | 0.98717 |  | 0.99608
1240 | 0.98977 | 0.98785 | 0.98709 |  | 0.99636
1250 | 0.98955 | 0.98775 | 0.98705 |  | 0.99668
1260 | 0.98932 | 0.98767 | 0.98704 |  | 0.99703
1270 | 0.9891 | 0.98759 | 0.98707 |  | 0.99739
1280 | 0.98887 | 0.98753 | 0.98713 |  | 0.99775
1290 | 0.98865 | 0.98748 | 0.98721 |  | 0.99808
1300 | 0.98842 | 0.98743 | 0.9873 |  | 0.99839
1310 | 0.98819 | 0.9874 | 0.9874 |  | 0.99867
1320 | 0.98797 | 0.98737 | 0.9875 |  | 0.99891
1330 | 0.98774 | 0.98735 | 0.9876 |  | 0.99911
1340 | 0.98752 | 0.98734 | 0.98771 |  | 0.99925
1350 | 0.9873 | 0.98734 | 0.98781 |  | 0.99933
1360 | 0.98708 | 0.98734 | 0.9879 |  | 0.99936
1370 | 0.98686 | 0.98735 | 0.988 |  | 0.99934
1380 | 0.98665 | 0.98736 | 0.98809 |  | 0.99927
1390 | 0.98644 | 0.98738 | 0.98818 |  | 0.99917
1400 | 0.98624 | 0.9874 | 0.98826 |  | 0.99903
1410 | 0.98604 | 0.98743 | 0.98834 |  | 0.99887
1420 | 0.98584 | 0.98746 | 0.98841 |  | 0.99868
1430 | 0.98565 | 0.9875 | 0.98848 |  | 0.99847
1440 | 0.98547 | 0.98754 | 0.98854 |  | 0.99826
1450 | 0.9853 | 0.98758 | 0.9886 |  | 0.99804
1460 | 0.98514 | 0.98763 | 0.98866 |  | 0.99781
1470 | 0.98498 | 0.98769 | 0.98871 |  | 0.9976
1480 | 0.98483 | 0.98775 | 0.98876 |  | 0.99739
1490 | 0.9847 | 0.98781 | 0.98881 |  | 0.99721
1500 | 0.98457 | 0.98788 | 0.98885 |  | 0.99705
1510 | 0.98446 | 0.98795 | 0.9889 |  | 0.99695
1520 | 0.98436 | 0.98802 | 0.98894 |  | 0.99689
1530 | 0.98428 | 0.9881 | 0.98899 |  | 0.99689
1540 | 0.98421 | 0.98819 | 0.98903 |  | 0.99693
1550 | 0.98415 | 0.98828 | 0.98908 |  | 0.997
1560 | 0.98411 | 0.98837 | 0.98913 |  | 0.99709
1570 | 0.98409 | 0.98847 | 0.98917 |  | 0.99721
1580 | 0.98409 | 0.98857 | 0.98922 |  | 0.99733
1590 | 0.9841 | 0.98868 | 0.98928 |  | 0.99748
1600 | 0.98413 | 0.98879 | 0.98934 |  | 0.99763
1610 | 0.98418 | 0.98891 | 0.98941 |  | 0.9978
1620 | 0.98425 | 0.98903 | 0.98948 |  | 0.99798
1630 | 0.98433 | 0.98915 | 0.98956 |  | 0.99816
1640 | 0.98444 | 0.98927 | 0.98965 |  | 0.99834
1650 | 0.98456 | 0.9894 | 0.98975 |  | 0.99851
1660 | 0.9847 | 0.98953 | 0.98985 |  | 0.99866
1670 | 0.98486 | 0.98966 | 0.98995 |  | 0.99877
1680 | 0.98503 | 0.98979 | 0.99006 |  | 0.99883
1690 | 0.98523 | 0.98993 | 0.99016 |  | 0.99883
1700 | 0.98544 | 0.99006 | 0.99027 |  | 0.99876
1710 | 0.98566 | 0.9902 | 0.99037 |  | 0.99861
1720 | 0.98591 | 0.99033 | 0.99047 |  | 0.9984
1730 | 0.98616 | 0.99046 | 0.99058 |  | 0.99813
1740 | 0.98643 | 0.99059 | 0.99068 |  | 0.99781
1750 | 0.9867 | 0.99072 | 0.99078 |  | 0.99744
1760 | 0.98698 | 0.99085 | 0.99088 |  | 0.99697
1770 | 0.98727 | 0.99097 | 0.99099 |  | 0.99634
1780 | 0.98755 | 0.99109 | 0.99109 |  | 0.99556
1790 | 0.98784 | 0.9912 | 0.99119 |  | 0.99462
1800 | 0.98813 | 0.99131 | 0.99129 |  | 0.99355

## Mirrors
Thorlabs Coatings
lbd [nm] | Protected Ag | Bare Gold
380 | 0.82198 | 0
390 | 0.8545 | 0
400 | 0.88718 | 0
410 | 0.9125 | 0
420 | 0.93099 | 0
430 | 0.94342 | 0
440 | 0.95124 | 0
450 | 0.95739 | 0
460 | 0.96078 | 0
470 | 0.96273 | 0
480 | 0.96437 | 0
490 | 0.96528 | 0.57
500 | 0.96552 | 0.6225
510 | 0.96504 | 0.6225
520 | 0.96423 | 0.66529
530 | 0.96372 | 0.73494
540 | 0.96334 | 0.78727
550 | 0.96341 | 0.82623
560 | 0.96218 | 0.85596
570 | 0.96072 | 0.87893
580 | 0.95904 | 0.89702
590 | 0.95853 | 0.91143
600 | 0.95754 | 0.92281
610 | 0.9573 | 0.9324
620 | 0.95721 | 0.9404
630 | 0.95713 | 0.9468
640 | 0.95712 | 0.95247
650 | 0.95723 | 0.95719
660 | 0.95683 | 0.96144
670 | 0.95662 | 0.96471
680 | 0.95588 | 0.96727
690 | 0.95572 | 0.96944
700 | 0.95394 | 0.9709
710 | 0.95205 | 0.9722
720 | 0.95091 | 0.97323
730 | 0.95141 | 0.97423
740 | 0.95171 | 0.97491
750 | 0.95261 | 0.97579
760 | 0.953 | 0.97617
770 | 0.95346 | 0.97706
780 | 0.95396 | 0.97758
790 | 0.95475 | 0.97811
800 | 0.95531 | 0.97867
810 | 0.95781 | 0.97927
820 | 0.95946 | 0.97916
830 | 0.95966 | 0.97908
840 | 0.95971 | 0.97968
850 | 0.96119 | 0.97995
860 | 0.96311 | 0.97983
870 | 0.96342 | 0.9803
880 | 0.96466 | 0.98058
890 | 0.96511 | 0.9807
900 | 0.96605 | 0.98172
910 | 0.96637 | 0.98058
920 | 0.96656 | 0.9808
930 | 0.96714 | 0.98132
940 | 0.96788 | 0.98094
950 | 0.96847 | 0.98104
960 | 0.96881 | 0.98098
970 | 0.96903 | 0.98111
980 | 0.96933 | 0.98109
990 | 0.9697 | 0.98114
1000 | 0.97014 | 0.98118
1010 | 0.97059 | 0.98134
1020 | 0.97098 | 0.98147
1030 | 0.97152 | 0.98144
1040 | 0.9721 | 0.98113
1050 | 0.97245 | 0.98152
1060 | 0.97267 | 0.98165
1070 | 0.9729 | 0.9815
1080 | 0.97339 | 0.98139
1090 | 0.9737 | 0.98149
1100 | 0.97401 | 0.9812
1110 | 0.97434 | 0.98124
1120 | 0.97456 | 0.98113
1130 | 0.97496 | 0.98124
1140 | 0.97538 | 0.9815
1150 | 0.97565 | 0.98162
1160 | 0.97579 | 0.98138
1170 | 0.9761 | 0.9811
1180 | 0.97632 | 0.98115
1190 | 0.97643 | 0.98133
1200 | 0.9768 | 0.98174
1210 | 0.97707 | 0.98099
1220 | 0.97708 | 0.98091
1230 | 0.97755 | 0.98076
1240 | 0.97797 | 0.98124
1250 | 0.97807 | 0.98135
1260 | 0.97834 | 0.98164
1270 | 0.97863 | 0.98145
1280 | 0.97885 | 0.98171
1290 | 0.9791 | 0.98155
1300 | 0.97917 | 0.98135
1310 | 0.97927 | 0.98167
1320 | 0.9796 | 0.98138
1330 | 0.97979 | 0.98142
1340 | 0.97966 | 0.98111
1350 | 0.97982 | 0.98089
1360 | 0.97985 | 0.98125
1370 | 0.98013 | 0.98204
1380 | 0.98049 | 0.98178
1390 | 0.98061 | 0.98189
1400 | 0.98069 | 0.98193
1410 | 0.98087 | 0.9822
1420 | 0.98098 | 0.98205
1430 | 0.98117 | 0.98175
1440 | 0.98112 | 0.9819
1450 | 0.98134 | 0.98207
1460 | 0.98167 | 0.98182
1470 | 0.98178 | 0.9823
1480 | 0.98202 | 0.98215
1490 | 0.98199 | 0.98191
1500 | 0.98173 | 0.98185
1510 | 0.98184 | 0.98168
1520 | 0.98193 | 0.98125
1530 | 0.98193 | 0.98114
1540 | 0.98214 | 0.98118
1550 | 0.98215 | 0.98142
1560 | 0.98192 | 0.9812
1570 | 0.9819 | 0.98108
1580 | 0.98212 | 0.98112
1590 | 0.98212 | 0.9811
1600 | 0.98237 | 0.98108
1610 | 0.98272 | 0.98149
1620 | 0.98292 | 0.98169
1630 | 0.983 | 0.98196
1640 | 0.98316 | 0.98214
1650 | 0.98319 | 0.98238
1660 | 0.98348 | 0.98264
1670 | 0.98349 | 0.98279
1680 | 0.98357 | 0.98217
1690 | 0.98386 | 0.98207
1700 | 0.98375 | 0.98212
1710 | 0.98392 | 0.98212
1720 | 0.98396 | 0.98161
1730 | 0.98353 | 0.98166
1740 | 0.98361 | 0.98147
1750 | 0.98333 | 0.98159
1760 | 0.98323 | 0.98133
1770 | 0.98313 | 0.98096
1780 | 0.98348 | 0.98125
1790 | 0.98341 | 0.98129
1800 | 0.98316 | 0.98085
1810 | 0.9833 | 0.98039
1820 | 0.98327 | 0.9805
1830 | 0.98342 | 0.98051
1840 | 0.98331 | 0.98051
1850 | 0.98339 | 0.98113
1860 | 0.98392 | 0.98109
1870 | 0.98382 | 0.98188
1880 | 0.98419 | 0.98211
1890 | 0.98422 | 0.98236
1900 | 0.98427 | 0.98259
1910 | 0.98454 | 0.98257
1920 | 0.98486 | 0.98303
1930 | 0.98506 | 0.98352
1940 | 0.98516 | 0.98371
1950 | 0.98533 | 0.98398
1960 | 0.98497 | 0.98321
1970 | 0.98516 | 0.98277
1980 | 0.98515 | 0.98274
1990 | 0.98558 | 0.98265
2000 | 0.98532 | 0.98271
2010 | 0.98551 | 0.98304
2020 | 0.98542 | 0.98287
2030 | 0.98585 | 0.98253
2040 | 0.98499 | 0.98241
2050 | 0.98479 | 0.98181
2060 | 0.9848 | 0.98133
2070 | 0.98477 | 0.98162
2080 | 0.98445 | 0.9809
2090 | 0.98431 | 0.98139
2100 | 0.98428 | 0.9807
2110 | 0.98402 | 0.98072
2120 | 0.98399 | 0.98035
2130 | 0.98373 | 0.97998
2140 | 0.98358 | 0.98031
2150 | 0.9834 | 0.98049
2160 | 0.98429 | 0.98049
2170 | 0.98421 | 0.98106
2180 | 0.98404 | 0.98182
2190 | 0.98442 | 0.98185
2200 | 0.98372 | 0.98169
2210 | 0.98438 | 0.98157
2220 | 0.98403 | 0.98149
2230 | 0.98348 | 0.98248
2240 | 0.98338 | 0.98251
2250 | 0.98455 | 0.98246
2260 | 0.98485 | 0.98198
2270 | 0.98497 | 0.9814
2280 | 0.98434 | 0.98292
2290 | 0.98408 | 0.98125
2300 | 0.98492 | 0.98309
2310 | 0.98571 | 0.98312
2320 | 0.9843 | 0.98375
2330 | 0.98455 | 0.98263
2340 | 0.98453 | 0.98428
2350 | 0.98462 | 0.98356
2360 | 0.98532 | 0.9822
2370 | 0.98428 | 0.98369
2380 | 0.9848 | 0.98483
2390 | 0.98379 | 0.9848
2400 | 0.98396 | 0.98267

## Coupling summary
0.7" = 25% best case |  |  |  | 0.9" = median case | [6 empty] | 1.2" = 75% best case
Parameters of AO simulations: |  |  | Coupling HAF |  |  |  | Coupling HEF
Average wind speed | 10m/s | Seeing | I=9 - HAF | I=10 - HAF | I=11 - HAF | I=12 - HAF | I= 9 - HEF | I=10 - HEF | I=11 - HEF | I=12 - HEF
Loop frequency | 1000 Hz | 0.7 | 0.75117 | 0.69179 | 0.52889 | 0.41997 | 0.83207 | 0.80973 | 0.795 | 0.73978
RON | 0.3 e- | 0.9 | 0.68963 | 0.63583 | 0.5004 | 0.31465 | 0.78493 | 0.77709 | 0.71179 | 0.66553
Delay | 2.5ms | 1.2 | 0.5945 | 0.52803 | 0.3986 | 0.21617 | 0.70818 | 0.67804 | 0.62316 | 0.54549
Amp. Noise F | 1.4
QE | 80%
Optical efficiency | 0.75 + 2*Al surfaces
|  | HEF to HAF SNR loss |  | 0.7
Gain of HAF over HEF
|  | Column1 | I=9 - HAF | I=10 - HAF | I=11 - HAF | I=12 - HAF
|  | 0.7 | 1.2897 | 1.2205 | 0.95038 | 0.811
|  | 0.9 | 1.2551 | 1.1689 | 1.0043 | 0.67541
|  | 1.2 | 1.1992 | 1.1125 | 0.91377 | 0.56613

## Coupling 0.7"
Fiber coupling for 0.7" seeing @500nm
Coupling obtained from complete simulation of AO+fiber modes.
| Coupling HAF |  |  |  | Coupling HEF
lbd [nm] | I=9 - HAF | I=10 - HAF | I=11 - HAF | I=12 - HAF | I= 9 - HEF | I=10 - HEF | I=11 - HEF | I=12 - HEF
980 | 0.67497 | 0.59587 | 0.45137 | 0.34153 | 0.80607 | 0.78577 | 0.78173 | 0.71651
990 | 0.67867 | 0.59903 | 0.45599 | 0.34614 | 0.807 | 0.78654 | 0.78229 | 0.71722
1000 | 0.68051 | 0.6022 | 0.45224 | 0.34135 | 0.80791 | 0.7873 | 0.78285 | 0.71792
1010 | 0.68234 | 0.60536 | 0.44849 | 0.33656 | 0.80881 | 0.78805 | 0.7834 | 0.71862
1020 | 0.68583 | 0.6094 | 0.45285 | 0.34094 | 0.80971 | 0.7888 | 0.78394 | 0.71932
1030 | 0.68933 | 0.61344 | 0.45721 | 0.34531 | 0.81059 | 0.78954 | 0.78447 | 0.72
1040 | 0.69264 | 0.61749 | 0.46141 | 0.34957 | 0.81146 | 0.79028 | 0.78499 | 0.72069
1050 | 0.69594 | 0.62153 | 0.46562 | 0.35383 | 0.81232 | 0.791 | 0.7855 | 0.72136
1060 | 0.6991 | 0.62557 | 0.46939 | 0.3576 | 0.81316 | 0.79173 | 0.786 | 0.72203
1070 | 0.70226 | 0.62875 | 0.47316 | 0.36136 | 0.814 | 0.79244 | 0.7865 | 0.7227
1080 | 0.70445 | 0.63193 | 0.47535 | 0.36326 | 0.81483 | 0.79315 | 0.78698 | 0.72336
1090 | 0.70664 | 0.63511 | 0.47754 | 0.36516 | 0.81565 | 0.79385 | 0.78746 | 0.72402
1100 | 0.7085 | 0.63829 | 0.47714 | 0.36558 | 0.81645 | 0.79454 | 0.78792 | 0.72467
1110 | 0.71036 | 0.64147 | 0.47675 | 0.366 | 0.81725 | 0.79523 | 0.78838 | 0.72532
1120 | 0.71188 | 0.64348 | 0.47701 | 0.36363 | 0.81803 | 0.79591 | 0.78883 | 0.72596
1130 | 0.71341 | 0.64548 | 0.47727 | 0.36127 | 0.8188 | 0.79658 | 0.78927 | 0.72659
1140 | 0.7158 | 0.64749 | 0.48148 | 0.36532 | 0.81957 | 0.79725 | 0.7897 | 0.72722
1150 | 0.71819 | 0.6495 | 0.4857 | 0.36937 | 0.82032 | 0.79791 | 0.79013 | 0.72785
1160 | 0.72053 | 0.65151 | 0.48979 | 0.37337 | 0.82106 | 0.79856 | 0.79054 | 0.72847
1170 | 0.72288 | 0.65465 | 0.49387 | 0.37738 | 0.82179 | 0.7992 | 0.79094 | 0.72908
1180 | 0.72533 | 0.65778 | 0.49789 | 0.38137 | 0.82251 | 0.79984 | 0.79134 | 0.72969
1190 | 0.72779 | 0.66092 | 0.5019 | 0.38536 | 0.82322 | 0.80048 | 0.79173 | 0.7303
1200 | 0.73043 | 0.66405 | 0.50585 | 0.38934 | 0.82392 | 0.8011 | 0.7921 | 0.73089
1210 | 0.73307 | 0.66719 | 0.50979 | 0.39332 | 0.82461 | 0.80172 | 0.79247 | 0.73149
1220 | 0.73568 | 0.67029 | 0.51354 | 0.39715 | 0.82528 | 0.80233 | 0.79283 | 0.73208
1230 | 0.7383 | 0.67339 | 0.51728 | 0.40099 | 0.82595 | 0.80293 | 0.79318 | 0.73266
1240 | 0.74071 | 0.67649 | 0.52055 | 0.40442 | 0.82661 | 0.80353 | 0.79352 | 0.73324
1250 | 0.74313 | 0.67959 | 0.52382 | 0.40784 | 0.82725 | 0.80412 | 0.79386 | 0.73381
1260 | 0.74396 | 0.6827 | 0.51735 | 0.40275 | 0.82788 | 0.80471 | 0.79418 | 0.73438
1270 | 0.7448 | 0.68398 | 0.51089 | 0.39765 | 0.82851 | 0.80529 | 0.79449 | 0.73494
1280 | 0.74648 | 0.68527 | 0.51272 | 0.39962 | 0.82912 | 0.80586 | 0.7948 | 0.7355
1290 | 0.74817 | 0.68656 | 0.51455 | 0.40159 | 0.82972 | 0.80642 | 0.7951 | 0.73605
1300 | 0.74902 | 0.68785 | 0.51309 | 0.39642 | 0.83031 | 0.80698 | 0.79538 | 0.73659
1310 | 0.74987 | 0.68914 | 0.51162 | 0.39123 | 0.83089 | 0.80753 | 0.79566 | 0.73714
1320 | 0.75176 | 0.69078 | 0.51514 | 0.39514 | 0.83146 | 0.80807 | 0.79593 | 0.73767
1330 | 0.75364 | 0.69243 | 0.51865 | 0.39905 | 0.83202 | 0.80861 | 0.79619 | 0.7382
1340 | 0.75547 | 0.69408 | 0.52208 | 0.40292 | 0.83257 | 0.80914 | 0.79645 | 0.73873
1350 | 0.75728 | 0.69572 | 0.52552 | 0.40679 | 0.83311 | 0.80966 | 0.79669 | 0.73925
1360 | 0.75911 | 0.69737 | 0.5289 | 0.41063 | 0.83363 | 0.81018 | 0.79692 | 0.73976
1370 | 0.76094 | 0.69968 | 0.53229 | 0.41448 | 0.83415 | 0.81069 | 0.79715 | 0.74027
1380 | 0.76273 | 0.70199 | 0.53562 | 0.41829 | 0.83466 | 0.81119 | 0.79736 | 0.74078
1390 | 0.76452 | 0.7043 | 0.53895 | 0.4221 | 0.83515 | 0.81168 | 0.79757 | 0.74128
1400 | 0.76625 | 0.70661 | 0.54223 | 0.42588 | 0.83563 | 0.81217 | 0.79777 | 0.74177
1410 | 0.76797 | 0.70893 | 0.5455 | 0.42965 | 0.83611 | 0.81266 | 0.79796 | 0.74226
1420 | 0.76965 | 0.71022 | 0.54873 | 0.43339 | 0.83657 | 0.81313 | 0.79814 | 0.74274
1430 | 0.77133 | 0.71151 | 0.55195 | 0.43713 | 0.83702 | 0.8136 | 0.79831 | 0.74322
1440 | 0.77167 | 0.7128 | 0.54727 | 0.43228 | 0.83746 | 0.81406 | 0.79847 | 0.7437
1450 | 0.77201 | 0.71409 | 0.54259 | 0.42743 | 0.83789 | 0.81452 | 0.79863 | 0.74416
1460 | 0.77356 | 0.71538 | 0.54569 | 0.43093 | 0.83831 | 0.81496 | 0.79877 | 0.74463
1470 | 0.77512 | 0.71723 | 0.5488 | 0.43444 | 0.83872 | 0.81541 | 0.7989 | 0.74508
1480 | 0.77649 | 0.71908 | 0.5517 | 0.43778 | 0.83912 | 0.81584 | 0.79903 | 0.74554
1490 | 0.77786 | 0.72092 | 0.55461 | 0.44113 | 0.83951 | 0.81627 | 0.79915 | 0.74598
1500 | 0.77889 | 0.72277 | 0.55716 | 0.44423 | 0.83988 | 0.81669 | 0.79926 | 0.74643
1510 | 0.77991 | 0.72462 | 0.5597 | 0.44733 | 0.84025 | 0.8171 | 0.79936 | 0.74686
1520 | 0.78031 | 0.7253 | 0.56158 | 0.45001 | 0.8406 | 0.81751 | 0.79945 | 0.74729
1530 | 0.78072 | 0.72599 | 0.56346 | 0.45269 | 0.84095 | 0.81791 | 0.79953 | 0.74772
1540 | 0.78047 | 0.72668 | 0.56441 | 0.45465 | 0.84128 | 0.8183 | 0.7996 | 0.74814
1550 | 0.78022 | 0.72737 | 0.56537 | 0.45661 | 0.8416 | 0.81869 | 0.79966 | 0.74856
1560 | 0.778 | 0.72806 | 0.55946 | 0.452 | 0.84192 | 0.81907 | 0.79972 | 0.74897
1570 | 0.77579 | 0.72789 | 0.55355 | 0.44739 | 0.84222 | 0.81944 | 0.79976 | 0.74937
1580 | 0.77645 | 0.72773 | 0.55621 | 0.45076 | 0.84251 | 0.81981 | 0.7998 | 0.74977
1590 | 0.77711 | 0.72756 | 0.55888 | 0.45413 | 0.84279 | 0.82017 | 0.79983 | 0.75017
1600 | 0.77772 | 0.7274 | 0.56145 | 0.45738 | 0.84306 | 0.82052 | 0.79985 | 0.75055
1610 | 0.77833 | 0.72724 | 0.56403 | 0.46064 | 0.84332 | 0.82086 | 0.79986 | 0.75094
1620 | 0.77889 | 0.72832 | 0.56653 | 0.46379 | 0.84356 | 0.8212 | 0.79986 | 0.75132
1630 | 0.77945 | 0.72941 | 0.56902 | 0.46694 | 0.8438 | 0.82154 | 0.79985 | 0.75169
1640 | 0.77993 | 0.7305 | 0.57143 | 0.46998 | 0.84403 | 0.82186 | 0.79983 | 0.75206
1650 | 0.78042 | 0.73158 | 0.57384 | 0.47302 | 0.84424 | 0.82218 | 0.79981 | 0.75242
1660 | 0.78086 | 0.73267 | 0.57617 | 0.47597 | 0.84445 | 0.82249 | 0.79977 | 0.75278
1670 | 0.7813 | 0.7336 | 0.5785 | 0.47891 | 0.84464 | 0.8228 | 0.79973 | 0.75313
1680 | 0.78174 | 0.73453 | 0.58077 | 0.48177 | 0.84483 | 0.82309 | 0.79967 | 0.75348
1690 | 0.78218 | 0.73546 | 0.58305 | 0.48463 | 0.845 | 0.82338 | 0.79961 | 0.75382
1700 | 0.78261 | 0.73638 | 0.58526 | 0.48741 | 0.84516 | 0.82367 | 0.79954 | 0.75416
1710 | 0.78304 | 0.73731 | 0.58748 | 0.49019 | 0.84531 | 0.82395 | 0.79946 | 0.75449
1720 | 0.78196 | 0.73713 | 0.57494 | 0.48344 | 0.84545 | 0.82422 | 0.79937 | 0.75482
1730 | 0.78088 | 0.73694 | 0.5624 | 0.47669 | 0.84558 | 0.82448 | 0.79927 | 0.75514
1740 | 0.78119 | 0.73676 | 0.56464 | 0.47919 | 0.8457 | 0.82474 | 0.79917 | 0.75546
1750 | 0.78151 | 0.73658 | 0.56687 | 0.48169 | 0.84581 | 0.82499 | 0.79905 | 0.75577
1760 | 0.78167 | 0.73639 | 0.56897 | 0.48406 | 0.84591 | 0.82523 | 0.79893 | 0.75608
1770 | 0.78183 | 0.73693 | 0.57106 | 0.48642 | 0.84599 | 0.82547 | 0.79879 | 0.75638
1780 | 0.78187 | 0.73746 | 0.573 | 0.48865 | 0.84607 | 0.8257 | 0.79865 | 0.75667
1790 | 0.7819 | 0.738 | 0.57495 | 0.49088 | 0.84614 | 0.82592 | 0.7985 | 0.75696
1800 | 0.78183 | 0.73853 | 0.57677 | 0.49297 | 0.84619 | 0.82614 | 0.79834 | 0.75724
| 0.75117 | 0.69179 | 0.52889 | 0.41997 | 0.83207 | 0.80973 | 0.795 | 0.73978

## Coupling 0.9"
Fiber coupling for 0.9" seeing @500nm
Coupling obtained from complete simulation of AO+fiber modes.
| Coupling HAF |  |  |  | Coupling HEF
lbd [nm] | I=9 - HAF | I=10 - HAF | I=11 - HAF | I=12 - HAF | I= 9 - HEF | I=10 - HEF | I=11 - HEF | I=12 - HEF
980 | 0.57847 | 0.52041 | 0.38182 | 0.23182 | 0.74036 | 0.74667 | 0.67214 | 0.62527
990 | 0.58251 | 0.52403 | 0.38712 | 0.23258 | 0.74182 | 0.74763 | 0.6733 | 0.62648
1000 | 0.58656 | 0.52765 | 0.38701 | 0.23335 | 0.74327 | 0.74858 | 0.67445 | 0.62769
1010 | 0.5906 | 0.53127 | 0.3869 | 0.23411 | 0.7447 | 0.74952 | 0.6756 | 0.62889
1020 | 0.59506 | 0.53606 | 0.39192 | 0.23775 | 0.74612 | 0.75046 | 0.67674 | 0.63009
1030 | 0.59951 | 0.54086 | 0.39694 | 0.24138 | 0.74752 | 0.75139 | 0.67787 | 0.63127
1040 | 0.60396 | 0.54565 | 0.40179 | 0.24502 | 0.74891 | 0.75231 | 0.679 | 0.63244
1050 | 0.60842 | 0.55044 | 0.40664 | 0.24865 | 0.75029 | 0.75322 | 0.68012 | 0.63361
1060 | 0.61287 | 0.55524 | 0.41109 | 0.25229 | 0.75165 | 0.75412 | 0.68123 | 0.63477
1070 | 0.61644 | 0.55903 | 0.41554 | 0.25493 | 0.753 | 0.75501 | 0.68234 | 0.63592
1080 | 0.62002 | 0.56282 | 0.41865 | 0.25758 | 0.75434 | 0.7559 | 0.68343 | 0.63706
1090 | 0.62359 | 0.56661 | 0.42176 | 0.26023 | 0.75566 | 0.75678 | 0.68452 | 0.63819
1100 | 0.62716 | 0.5704 | 0.42372 | 0.26287 | 0.75696 | 0.75765 | 0.6856 | 0.63931
1110 | 0.63074 | 0.57419 | 0.42568 | 0.26552 | 0.75826 | 0.75852 | 0.68668 | 0.64042
1120 | 0.63334 | 0.57636 | 0.42706 | 0.26544 | 0.75954 | 0.75937 | 0.68775 | 0.64153
1130 | 0.63593 | 0.57853 | 0.42843 | 0.26537 | 0.76081 | 0.76022 | 0.68881 | 0.64262
1140 | 0.63853 | 0.58069 | 0.4329 | 0.26529 | 0.76206 | 0.76106 | 0.68986 | 0.64371
1150 | 0.64113 | 0.58286 | 0.43737 | 0.26522 | 0.7633 | 0.76189 | 0.69091 | 0.64479
1160 | 0.64373 | 0.58503 | 0.44178 | 0.26514 | 0.76452 | 0.76271 | 0.69195 | 0.64585
1170 | 0.64705 | 0.58873 | 0.44619 | 0.26897 | 0.76573 | 0.76352 | 0.69298 | 0.64692
1180 | 0.65038 | 0.59243 | 0.45062 | 0.27279 | 0.76693 | 0.76433 | 0.694 | 0.64797
1190 | 0.6537 | 0.59613 | 0.45505 | 0.27661 | 0.76811 | 0.76513 | 0.69502 | 0.64901
1200 | 0.65702 | 0.59983 | 0.4595 | 0.28043 | 0.76928 | 0.76592 | 0.69603 | 0.65005
1210 | 0.66035 | 0.60353 | 0.46395 | 0.28425 | 0.77044 | 0.7667 | 0.69704 | 0.65107
1220 | 0.66365 | 0.60705 | 0.46829 | 0.2878 | 0.77158 | 0.76748 | 0.69803 | 0.65209
1230 | 0.66696 | 0.61057 | 0.47263 | 0.29134 | 0.77271 | 0.76824 | 0.69902 | 0.6531
1240 | 0.67027 | 0.6141 | 0.47659 | 0.29489 | 0.77383 | 0.769 | 0.7 | 0.6541
1250 | 0.67358 | 0.61762 | 0.48055 | 0.29843 | 0.77493 | 0.76975 | 0.70098 | 0.65509
1260 | 0.67688 | 0.62114 | 0.47928 | 0.30197 | 0.77601 | 0.77049 | 0.70194 | 0.65607
1270 | 0.67887 | 0.62275 | 0.47801 | 0.30166 | 0.77709 | 0.77123 | 0.7029 | 0.65704
1280 | 0.68086 | 0.62435 | 0.48045 | 0.30135 | 0.77815 | 0.77195 | 0.70386 | 0.65801
1290 | 0.68285 | 0.62596 | 0.48289 | 0.30104 | 0.77919 | 0.77267 | 0.7048 | 0.65896
1300 | 0.68484 | 0.62757 | 0.48285 | 0.30073 | 0.78022 | 0.77338 | 0.70574 | 0.65991
1310 | 0.68683 | 0.62917 | 0.48281 | 0.30041 | 0.78124 | 0.77408 | 0.70667 | 0.66085
1320 | 0.68888 | 0.63126 | 0.48663 | 0.3007 | 0.78225 | 0.77477 | 0.70759 | 0.66178
1330 | 0.69093 | 0.63334 | 0.49044 | 0.30099 | 0.78324 | 0.77546 | 0.70851 | 0.6627
1340 | 0.69298 | 0.63543 | 0.49416 | 0.30128 | 0.78421 | 0.77614 | 0.70942 | 0.66361
1350 | 0.69503 | 0.63752 | 0.49789 | 0.30157 | 0.78518 | 0.77681 | 0.71032 | 0.66451
1360 | 0.69708 | 0.6396 | 0.50157 | 0.30186 | 0.78613 | 0.77747 | 0.71122 | 0.66541
1370 | 0.69951 | 0.64244 | 0.50525 | 0.30545 | 0.78706 | 0.77812 | 0.7121 | 0.66629
1380 | 0.70194 | 0.64527 | 0.50885 | 0.30904 | 0.78798 | 0.77877 | 0.71298 | 0.66717
1390 | 0.70436 | 0.64811 | 0.51246 | 0.31263 | 0.78889 | 0.7794 | 0.71386 | 0.66804
1400 | 0.70679 | 0.65094 | 0.51597 | 0.31622 | 0.78978 | 0.78003 | 0.71472 | 0.6689
1410 | 0.70922 | 0.65377 | 0.51949 | 0.31981 | 0.79066 | 0.78065 | 0.71558 | 0.66975
1420 | 0.71092 | 0.65526 | 0.52293 | 0.32025 | 0.79153 | 0.78127 | 0.71644 | 0.67059
1430 | 0.71262 | 0.65675 | 0.52637 | 0.32069 | 0.79239 | 0.78187 | 0.71728 | 0.67142
1440 | 0.71433 | 0.65824 | 0.52443 | 0.32112 | 0.79322 | 0.78247 | 0.71812 | 0.67225
1450 | 0.71603 | 0.65972 | 0.52248 | 0.32156 | 0.79405 | 0.78306 | 0.71895 | 0.67306
1460 | 0.71773 | 0.66121 | 0.52575 | 0.32199 | 0.79486 | 0.78364 | 0.71977 | 0.67387
1470 | 0.71967 | 0.66366 | 0.52903 | 0.32525 | 0.79566 | 0.78421 | 0.72058 | 0.67467
1480 | 0.72161 | 0.6661 | 0.5321 | 0.32851 | 0.79644 | 0.78478 | 0.72139 | 0.67546
1490 | 0.72354 | 0.66855 | 0.53517 | 0.33176 | 0.79721 | 0.78533 | 0.72219 | 0.67624
1500 | 0.72548 | 0.671 | 0.53791 | 0.33502 | 0.79797 | 0.78588 | 0.72299 | 0.67701
1510 | 0.72742 | 0.67345 | 0.54064 | 0.33828 | 0.79871 | 0.78642 | 0.72377 | 0.67777
1520 | 0.72824 | 0.67483 | 0.54279 | 0.34066 | 0.79944 | 0.78695 | 0.72455 | 0.67853
1530 | 0.72906 | 0.67621 | 0.54494 | 0.34304 | 0.80015 | 0.78748 | 0.72532 | 0.67927
1540 | 0.72989 | 0.67759 | 0.54623 | 0.34542 | 0.80085 | 0.78799 | 0.72609 | 0.68001
1550 | 0.73071 | 0.67897 | 0.54753 | 0.3478 | 0.80154 | 0.7885 | 0.72685 | 0.68074
1560 | 0.73153 | 0.68035 | 0.5448 | 0.35019 | 0.80222 | 0.789 | 0.7276 | 0.68146
1570 | 0.73159 | 0.68062 | 0.54208 | 0.35054 | 0.80287 | 0.78949 | 0.72834 | 0.68217
1580 | 0.73165 | 0.6809 | 0.5449 | 0.3509 | 0.80352 | 0.78998 | 0.72908 | 0.68287
1590 | 0.73171 | 0.68118 | 0.54772 | 0.35125 | 0.80415 | 0.79045 | 0.7298 | 0.68356
1600 | 0.73176 | 0.68145 | 0.55045 | 0.3516 | 0.80477 | 0.79092 | 0.73053 | 0.68425
1610 | 0.73182 | 0.68173 | 0.55317 | 0.35196 | 0.80537 | 0.79138 | 0.73124 | 0.68492
1620 | 0.73284 | 0.68329 | 0.55581 | 0.35534 | 0.80597 | 0.79183 | 0.73195 | 0.68559
1630 | 0.73385 | 0.68485 | 0.55844 | 0.35872 | 0.80654 | 0.79228 | 0.73265 | 0.68625
1640 | 0.73487 | 0.68641 | 0.56097 | 0.36209 | 0.80711 | 0.79271 | 0.73334 | 0.6869
1650 | 0.73588 | 0.68798 | 0.5635 | 0.36547 | 0.80766 | 0.79314 | 0.73403 | 0.68754
1660 | 0.7369 | 0.68954 | 0.56594 | 0.36885 | 0.80819 | 0.79356 | 0.7347 | 0.68817
1670 | 0.73777 | 0.6909 | 0.56838 | 0.37206 | 0.80871 | 0.79397 | 0.73537 | 0.68879
1680 | 0.73864 | 0.69227 | 0.57076 | 0.37527 | 0.80922 | 0.79437 | 0.73604 | 0.68941
1690 | 0.7395 | 0.69364 | 0.57315 | 0.37848 | 0.80971 | 0.79477 | 0.73669 | 0.69001
1700 | 0.74038 | 0.695 | 0.57546 | 0.38168 | 0.81019 | 0.79516 | 0.73734 | 0.69061
1710 | 0.74124 | 0.69637 | 0.57778 | 0.38489 | 0.81066 | 0.79554 | 0.73799 | 0.6912
1720 | 0.74138 | 0.69644 | 0.5735 | 0.38368 | 0.81112 | 0.79591 | 0.73862 | 0.69178
1730 | 0.74153 | 0.69652 | 0.56923 | 0.38248 | 0.81155 | 0.79627 | 0.73925 | 0.69235
1740 | 0.74167 | 0.69659 | 0.57141 | 0.38127 | 0.81198 | 0.79663 | 0.73987 | 0.69291
1750 | 0.74181 | 0.69666 | 0.57359 | 0.38006 | 0.81239 | 0.79697 | 0.74048 | 0.69346
1760 | 0.74195 | 0.69673 | 0.57561 | 0.37886 | 0.81279 | 0.79731 | 0.74109 | 0.69401
1770 | 0.74243 | 0.69768 | 0.57764 | 0.38155 | 0.81317 | 0.79764 | 0.74169 | 0.69454
1780 | 0.7429 | 0.69861 | 0.57953 | 0.38425 | 0.81354 | 0.79796 | 0.74228 | 0.69507
1790 | 0.74337 | 0.69955 | 0.58142 | 0.38694 | 0.8139 | 0.79828 | 0.74286 | 0.69559
1800 | 0.74385 | 0.70049 | 0.58319 | 0.38964 | 0.81424 | 0.79858 | 0.74344 | 0.6961
| 0.68963 | 0.63583 | 0.5004 | 0.31465 | 0.78493 | 0.77709 | 0.71179 | 0.66553

## Coupling 1.2"
Fiber coupling for 1.2" seeing @500nm
Coupling obtained from complete simulation of AO+fiber modes.
| Coupling HAF |  |  |  | Coupling HEF
lbd [nm] | I=9 - HAF | I=10 - HAF | I=11 - HAF | I=12 - HAF | I= 9 - HEF | I=10 - HEF | I=11 - HEF | I=12 - HEF
980 | 0.44823 | 0.37747 | 0.27043 | 0.14509 | 0.6385 | 0.61151 | 0.55996 | 0.48607
990 | 0.45305 | 0.38189 | 0.27544 | 0.14572 | 0.64064 | 0.61346 | 0.56177 | 0.48779
1000 | 0.45786 | 0.38631 | 0.27577 | 0.14635 | 0.64276 | 0.6154 | 0.56357 | 0.4895
1010 | 0.46268 | 0.39072 | 0.2761 | 0.14698 | 0.64487 | 0.61732 | 0.56536 | 0.49121
1020 | 0.46809 | 0.3962 | 0.28093 | 0.14975 | 0.64697 | 0.61923 | 0.56713 | 0.4929
1030 | 0.47351 | 0.40168 | 0.28576 | 0.15253 | 0.64904 | 0.62113 | 0.5689 | 0.49458
1040 | 0.47892 | 0.40716 | 0.29048 | 0.15531 | 0.6511 | 0.62301 | 0.57066 | 0.49625
1050 | 0.48434 | 0.41264 | 0.29521 | 0.15809 | 0.65314 | 0.62489 | 0.57241 | 0.49791
1060 | 0.48975 | 0.41811 | 0.29961 | 0.16086 | 0.65517 | 0.62675 | 0.57415 | 0.49956
1070 | 0.49435 | 0.42273 | 0.30401 | 0.16297 | 0.65718 | 0.6286 | 0.57588 | 0.5012
1080 | 0.49894 | 0.42734 | 0.30718 | 0.16509 | 0.65918 | 0.63044 | 0.57759 | 0.50284
1090 | 0.50353 | 0.43195 | 0.31036 | 0.1672 | 0.66116 | 0.63226 | 0.5793 | 0.50446
1100 | 0.50813 | 0.43656 | 0.3124 | 0.16931 | 0.66312 | 0.63408 | 0.581 | 0.50607
1110 | 0.51272 | 0.44117 | 0.31443 | 0.17142 | 0.66506 | 0.63588 | 0.58269 | 0.50767
1120 | 0.51625 | 0.44449 | 0.31583 | 0.17174 | 0.66699 | 0.63767 | 0.58437 | 0.50926
1130 | 0.51978 | 0.44781 | 0.31722 | 0.17207 | 0.6689 | 0.63945 | 0.58604 | 0.51084
1140 | 0.52331 | 0.45113 | 0.3218 | 0.17239 | 0.6708 | 0.64121 | 0.5877 | 0.51241
1150 | 0.52685 | 0.45445 | 0.32638 | 0.17272 | 0.67268 | 0.64297 | 0.58935 | 0.51397
1160 | 0.53038 | 0.45777 | 0.33095 | 0.17305 | 0.67454 | 0.64471 | 0.59099 | 0.51552
1170 | 0.53475 | 0.4625 | 0.33551 | 0.17613 | 0.67639 | 0.64644 | 0.59262 | 0.51706
1180 | 0.53912 | 0.46722 | 0.3401 | 0.17921 | 0.67822 | 0.64816 | 0.59424 | 0.51859
1190 | 0.54349 | 0.47195 | 0.34469 | 0.18229 | 0.68004 | 0.64987 | 0.59585 | 0.52011
1200 | 0.54786 | 0.47668 | 0.3493 | 0.18537 | 0.68183 | 0.65156 | 0.59745 | 0.52162
1210 | 0.55223 | 0.4814 | 0.35392 | 0.18845 | 0.68362 | 0.65324 | 0.59904 | 0.52312
1220 | 0.55647 | 0.48594 | 0.35845 | 0.1914 | 0.68538 | 0.65492 | 0.60062 | 0.52461
1230 | 0.56071 | 0.49049 | 0.36299 | 0.19434 | 0.68713 | 0.65657 | 0.60219 | 0.52609
1240 | 0.56494 | 0.49503 | 0.36722 | 0.19728 | 0.68886 | 0.65822 | 0.60375 | 0.52756
1250 | 0.56918 | 0.49957 | 0.37145 | 0.20022 | 0.69058 | 0.65985 | 0.6053 | 0.52902
1260 | 0.57342 | 0.50411 | 0.37068 | 0.20317 | 0.69228 | 0.66148 | 0.60685 | 0.53047
1270 | 0.57604 | 0.50653 | 0.3699 | 0.20304 | 0.69396 | 0.66309 | 0.60838 | 0.53191
1280 | 0.57867 | 0.50894 | 0.37282 | 0.20291 | 0.69563 | 0.66469 | 0.6099 | 0.53334
1290 | 0.58129 | 0.51135 | 0.37574 | 0.20278 | 0.69728 | 0.66627 | 0.61141 | 0.53475
1300 | 0.58392 | 0.51376 | 0.37568 | 0.20265 | 0.69891 | 0.66785 | 0.61291 | 0.53616
1310 | 0.58654 | 0.51617 | 0.37562 | 0.20253 | 0.70053 | 0.66941 | 0.6144 | 0.53756
1320 | 0.58936 | 0.51914 | 0.37976 | 0.20326 | 0.70213 | 0.67096 | 0.61589 | 0.53895
1330 | 0.59218 | 0.5221 | 0.3839 | 0.20399 | 0.70372 | 0.6725 | 0.61736 | 0.54033
1340 | 0.595 | 0.52506 | 0.38798 | 0.20472 | 0.70529 | 0.67403 | 0.61882 | 0.5417
1350 | 0.59782 | 0.52802 | 0.39206 | 0.20544 | 0.70684 | 0.67554 | 0.62028 | 0.54305
1360 | 0.60064 | 0.53099 | 0.3961 | 0.20617 | 0.70837 | 0.67705 | 0.62172 | 0.5444
1370 | 0.60391 | 0.53469 | 0.40014 | 0.20926 | 0.70989 | 0.67854 | 0.62315 | 0.54574
1380 | 0.60718 | 0.53839 | 0.40412 | 0.21234 | 0.7114 | 0.68002 | 0.62457 | 0.54707
1390 | 0.61046 | 0.5421 | 0.40809 | 0.21543 | 0.71288 | 0.68148 | 0.62599 | 0.54839
1400 | 0.61373 | 0.5458 | 0.412 | 0.21851 | 0.71435 | 0.68294 | 0.62739 | 0.5497
1410 | 0.617 | 0.54951 | 0.41591 | 0.22159 | 0.71581 | 0.68438 | 0.62879 | 0.55099
1420 | 0.61938 | 0.55192 | 0.41975 | 0.2216 | 0.71725 | 0.68581 | 0.63017 | 0.55228
1430 | 0.62176 | 0.55433 | 0.4236 | 0.22161 | 0.71867 | 0.68723 | 0.63154 | 0.55356
1440 | 0.62413 | 0.55675 | 0.4219 | 0.22162 | 0.72007 | 0.68864 | 0.63291 | 0.55483
1450 | 0.62651 | 0.55916 | 0.42019 | 0.22163 | 0.72146 | 0.69003 | 0.63426 | 0.55608
1460 | 0.62888 | 0.56158 | 0.42388 | 0.22164 | 0.72283 | 0.69142 | 0.63561 | 0.55733
1470 | 0.6316 | 0.56469 | 0.42757 | 0.22448 | 0.72419 | 0.69279 | 0.63694 | 0.55857
1480 | 0.63431 | 0.5678 | 0.43109 | 0.22733 | 0.72553 | 0.69415 | 0.63826 | 0.55979
1490 | 0.63702 | 0.57091 | 0.43461 | 0.23018 | 0.72685 | 0.69549 | 0.63958 | 0.56101
1500 | 0.63973 | 0.57402 | 0.43786 | 0.23303 | 0.72816 | 0.69683 | 0.64089 | 0.56222
1510 | 0.64244 | 0.57713 | 0.44111 | 0.23588 | 0.72945 | 0.69815 | 0.64218 | 0.56341
1520 | 0.64403 | 0.57915 | 0.44389 | 0.23805 | 0.73072 | 0.69946 | 0.64347 | 0.5646
1530 | 0.64562 | 0.58116 | 0.44667 | 0.24023 | 0.73198 | 0.70076 | 0.64474 | 0.56578
1540 | 0.64721 | 0.58317 | 0.44872 | 0.2424 | 0.73322 | 0.70205 | 0.64601 | 0.56694
1550 | 0.64881 | 0.58519 | 0.45077 | 0.24458 | 0.73445 | 0.70332 | 0.64726 | 0.5681
1560 | 0.6504 | 0.5872 | 0.44858 | 0.24675 | 0.73565 | 0.70459 | 0.64851 | 0.56925
1570 | 0.65118 | 0.58822 | 0.44638 | 0.24696 | 0.73685 | 0.70584 | 0.64975 | 0.57038
1580 | 0.65197 | 0.58923 | 0.44977 | 0.24716 | 0.73802 | 0.70708 | 0.65097 | 0.57151
1590 | 0.65275 | 0.59024 | 0.45315 | 0.24736 | 0.73918 | 0.7083 | 0.65219 | 0.57262
1600 | 0.65353 | 0.59125 | 0.45644 | 0.24757 | 0.74033 | 0.70952 | 0.65339 | 0.57373
1610 | 0.65432 | 0.59227 | 0.45973 | 0.24777 | 0.74145 | 0.71072 | 0.65459 | 0.57482
1620 | 0.65609 | 0.59457 | 0.46293 | 0.25093 | 0.74256 | 0.71191 | 0.65578 | 0.57591
1630 | 0.65787 | 0.59687 | 0.46614 | 0.25409 | 0.74365 | 0.71309 | 0.65695 | 0.57698
1640 | 0.65965 | 0.59917 | 0.46924 | 0.25724 | 0.74473 | 0.71426 | 0.65812 | 0.57805
1650 | 0.66142 | 0.60147 | 0.47235 | 0.2604 | 0.74579 | 0.71542 | 0.65928 | 0.5791
1660 | 0.6632 | 0.60377 | 0.47537 | 0.26356 | 0.74684 | 0.71656 | 0.66043 | 0.58015
1670 | 0.66479 | 0.60586 | 0.47839 | 0.26665 | 0.74787 | 0.71769 | 0.66156 | 0.58118
1680 | 0.66638 | 0.60795 | 0.48135 | 0.26975 | 0.74888 | 0.71881 | 0.66269 | 0.58221
1690 | 0.66797 | 0.61004 | 0.48431 | 0.27284 | 0.74987 | 0.71992 | 0.66381 | 0.58322
1700 | 0.66956 | 0.61213 | 0.4872 | 0.27594 | 0.75085 | 0.72101 | 0.66492 | 0.58423
1710 | 0.67115 | 0.61422 | 0.49009 | 0.27904 | 0.75182 | 0.7221 | 0.66602 | 0.58522
1720 | 0.67177 | 0.61468 | 0.48609 | 0.27829 | 0.75276 | 0.72317 | 0.6671 | 0.58621
1730 | 0.67239 | 0.61514 | 0.4821 | 0.27755 | 0.75369 | 0.72423 | 0.66818 | 0.58718
1740 | 0.673 | 0.6156 | 0.48482 | 0.27681 | 0.75461 | 0.72528 | 0.66925 | 0.58815
1750 | 0.67362 | 0.61606 | 0.48754 | 0.27607 | 0.7555 | 0.72631 | 0.67031 | 0.5891
1760 | 0.67423 | 0.61652 | 0.49013 | 0.27532 | 0.75638 | 0.72734 | 0.67136 | 0.59005
1770 | 0.67537 | 0.61815 | 0.49271 | 0.27807 | 0.75725 | 0.72835 | 0.6724 | 0.59098
1780 | 0.67651 | 0.61977 | 0.49517 | 0.28082 | 0.7581 | 0.72935 | 0.67343 | 0.5919
1790 | 0.67765 | 0.6214 | 0.49763 | 0.28357 | 0.75893 | 0.73033 | 0.67445 | 0.59282
1800 | 0.67879 | 0.62303 | 0.49997 | 0.28632 | 0.75974 | 0.73131 | 0.67546 | 0.59372
AVG | 0.5945 | 0.52803 | 0.3986 | 0.21617 | 0.70818 | 0.67804 | 0.62316 | 0.54549

## All, minus coupling
Dummy table for ipython notebook "transmission vs fiber size"
Content of the table: transmission of all components, except fiber coupling
lbd [nm] | Throughut HAF | Throughput HEF
980 | 0.061779 | 0.059058
990 | 0.061779 | 0.059058
1000 | 0.063972 | 0.061168
1010 | 0.065717 | 0.062851
1020 | 0.06796 | 0.065021
1030 | 0.070508 | 0.067473
1040 | 0.073593 | 0.070428
1050 | 0.076742 | 0.073456
1060 | 0.079886 | 0.07648
1070 | 0.083071 | 0.079517
1080 | 0.084401 | 0.080806
1090 | 0.088744 | 0.08495
1100 | 0.091166 | 0.087284
1110 | 0.093946 | 0.08993
1120 | 0.095943 | 0.091826
1130 | 0.097346 | 0.093151
1140 | 0.096101 | 0.091959
1150 | 0.099451 | 0.095146
1160 | 0.10036 | 0.096018
1170 | 0.10124 | 0.096816
1180 | 0.10205 | 0.097573
1190 | 0.10189 | 0.097389
1200 | 0.10249 | 0.097936
1210 | 0.10348 | 0.09885
1220 | 0.10365 | 0.098973
1230 | 0.10473 | 0.099982
1240 | 0.10536 | 0.10056
1250 | 0.10595 | 0.10108
1260 | 0.10683 | 0.10186
1270 | 0.10798 | 0.10291
1280 | 0.10925 | 0.10409
1290 | 0.11067 | 0.10541
1300 | 0.11163 | 0.10627
1310 | 0.11216 | 0.10672
1320 | 0.11316 | 0.10763
1330 | 0.11396 | 0.10834
1340 | 0.11471 | 0.109
1350 | 0.115 | 0.10922
1360 | 0.11538 | 0.10953
1370 | 0.11553 | 0.10962
1380 | 0.11481 | 0.10889
1390 | 0.11453 | 0.10856
1400 | 0.11469 | 0.10866
1410 | 0.11432 | 0.10827
1420 | 0.11389 | 0.10781
1430 | 0.11266 | 0.1066
1440 | 0.11409 | 0.10788
1450 | 0.11461 | 0.10834
1460 | 0.11458 | 0.10825
1470 | 0.11345 | 0.10715
1480 | 0.11334 | 0.10697
1490 | 0.1131 | 0.10672
1500 | 0.11225 | 0.10587
1510 | 0.11233 | 0.10589
1520 | 0.11268 | 0.10618
1530 | 0.11357 | 0.10698
1540 | 0.11418 | 0.10751
1550 | 0.11478 | 0.10804
1560 | 0.11363 | 0.10692
1570 | 0.11502 | 0.10821
1580 | 0.11629 | 0.10937
1590 | 0.11716 | 0.11016
1600 | 0.11743 | 0.11038
1610 | 0.11721 | 0.11017
1620 | 0.11862 | 0.11146
1630 | 0.11924 | 0.11203
1640 | 0.11989 | 0.11264
1650 | 0.12027 | 0.11299
1660 | 0.12073 | 0.11341
1670 | 0.12083 | 0.11351
1680 | 0.12041 | 0.11312
1690 | 0.12151 | 0.11415
1700 | 0.1217 | 0.11436
1710 | 0.12108 | 0.11379
1720 | 0.12 | 0.11277
1730 | 0.11815 | 0.11107
1740 | 0.11757 | 0.11057
1750 | 0.11702 | 0.11006
1760 | 0.11594 | 0.10911
1770 | 0.11407 | 0.10739
1780 | 0.11198 | 0.10546
1790 | 0.10945 | 0.10314
1800 | 0.1059 | 0.09985